Clarification to BID doc of Procurement of Medical and Diagnostic equipment and Medical Furniture for "Center of Hematology named after R.H. Yeolyan" ICB No: CR4/ICB/B-G/012-15

1. Question: Item 8 - Oxygen Concentrator
1) With ref to the spec required:

Compact and easy to transport oxygen concentrator (Mobile on Castors)

Could you please specify if a PORTABLE OXYGEN CONCENTRATOR or a TRANSPORTABLE OXYGEN CONCENTRATOR is required?

2) With ref to technical specification:

Flow volume: 6 l/min –

Could you confirm if is required:

Concentrator with Continuous Flow (stationary concentrator) 

Concentrator with Continuous Flow/On demand: concentrator transportable 

Concentrator with Flow on demand: portable concentrator 

Please note that:

The Transportable Concentrator arrives up to maximum of 3 l/m with continuous flow and up to a maximum of 10 l/min on demand, while the portable concentrator arrives up to a maximum of  6 l/min on demand 

3) With ref to technical specification:

Sound Level (ANSI): not more than 35dB

Please consider that on the market there is only one model that respect the limit of 35 db, which is portable (arrives up to flow 6l/m on demand).  Therefore, should we assume that a portable concentrator is required?

4) With ref to technical specification:

Steps: what do you mean with Steps as follows:

Flow:

0...2 lpm in 0.1 steps

2...4 lpm in 0.2 steps

4...6 lpm in 0.5 steps

Please note that both the Transportable and the Portable concentrators increase of 1 l/m at the time, increments of 0,5 l/min o 0,2 l/min are not possible on these type of concentrators, but only on Stationary concentrators.

Please clarify

5) With ref to technical specification:

Accessories: Humidifier, nasal canula: 2 pcs.

Transportable and the Portable concentrators do not require Humidification.

Could you please clarify this.

Answer:  
1) In the following case, we mean a reliable oxygen concentrator for both domestic and clinical use, which would be optimized for the production of oxygen at home.The mentioned electronically controlled oxygen concentrator separates oxygen from the ambient air and delivers it to the user by nasal cannula.
2) The proposed device should haveup to 6 l/min oxygen flow adjustment feature (the maximum recommended flow should be 6 l/min), and the oxygen concentration should not be less than 75% in case of 6 l/min. Theproposeddevice must meet the MGP classification (Law on Medical Devices) II a class.
3) The Oxygen Concentrator is used for long-term oxygen therapy, and taking into account that in the majority of cases continuous mode of operation of the device is required, special attention should be paid to the noise level, as each extra decibel acquires significance because the increased noisiness during hours creates unfavorable emotional background which will lead to irritation. The human ear perceives the noise level of 20…25 dB, but unfortunately oxygen concentrator with less than 20 dB is not available yet, therefore the most silent oxygen concentrator is considered the concentrator with  the noise of 31…35 dB.In the medical devices market most of high end class concentrators pass under this position.
4) please be assure taht, we have checked the technical specifications indicated in the tender and we would like to inform you that most of the high end class concentrators pass under the mentioned technical specifications.
5) please be assure that, we have checked the technical specifications indicated in the tender and we would like to inform you thatthepackageof the proposed device should include: corner connecter of humidifier, reusable humidifier, coarse filter and nasal cannula.

2. Question: Item 16 - UV-Air Flow Cleaner-Recirculator
The following technical specifications are required:

Performance m3/h: not less than 100

The bactericidal efficacy: not less than 99%

Could you please specify:

The real data for performance and bactericidal efficacyrequired, as the data indicated above, from our knowledge and from market’s research, seem to be too high.

Answer: please be assure that we have checked the technical specifications indicated in the tender and we would like to inform you that most of the UV-AirFlowCleaner-Recirculators are acceptable under the mentioned technical specifications.

