# PROCUREMENT DOCUMENTS

Bidding Document for

Procurement of Goods

Procurement of:

Medical, diagnostic equipment, Medical supplies and furniture for “Vanadzor MC” CJSC, comprising of 4 lots

Lot 1- Supply and Installation of Diagnostic and Imaging equipment for "Vanadzor MC” CJSC

Lot 2- Supply and Installation of Medical Laboratory equipment for "Vanadzor MC” CJSC

Lot 3 - Supply and Installation of Surgery and Hospital equipment for "Vanadzor MC” CJSC

Lot 4 -Supply and Installation of Medical Furniture for "Vanadzor MC” CJSC

ICB No: CR4/ICB/B-G/013-16

Project: Disease Prevention and Control

Project, Credit No. 5222-AM

Purchaser: “Health projects implementation unit” SA

Country: Republic of Armenia

Issued on: 16 January, 2017

Table of Contents

PART 1 – Bidding Procedures	1

Section I.  Instructions to Bidders	3

Section II.  Bid Data Sheet (BDS)	29

Section III.  Evaluation and Qualification Criteria	37

Section IV.  Bidding Forms	41

Section V.  Eligible Countries	57

Section VI. Bank Policy - Corrupt and Fraudulent Practices	59

PART 2 – Supply Requirements	63

Section VII.  Schedule of Requirements	64

PART 3 - Contract	185

Section VIII.  General Conditions of Contract	187

Section IX.  Special Conditions of Contract	211

Section X.  Contract Forms	219

# PART 1 – Bidding Procedures

Section I. Instructions to Bidders

Table of Clauses

A.	General	5

1.	Scope of Bid	5

2.	Source of Funds	5

3.	Corrupt and Fraudulent Practices	5

4.	Eligible Bidders	6

5.	Eligible Goods and Related Services	8

B. Contents of Bidding Document	9

6.	Sections of Bidding Document	9

7.	Clarification of Bidding Documents, Site Visit, Pre-Bid Meeting	10

8.	Amendment of Bidding Document	10

C. Preparation of Bids	10

9.	Cost of Bidding	10

10.	Language of Bid	10

11.	Documents Comprising the Bid	10

12.	Letter of Bid and Price Schedules	11

13.	Alternative Bids	11

14.	Bid Prices and Discounts	12

15.	Currencies of Bid and Payment	14

16.	Documents Establishing the Eligibility and Conformity of the Goods and Related Services	14

17.	Documents Establishing the Eligibility and Qualifications of  the Bidder	15

18.	Period of Validity of Bids	15

19.	Bid Security	16

20.	Format and Signing of Bid	18

D. Submission and Opening of Bids	18

21.	Sealing and Marking of Bids	18

22.	Deadline for Submission of Bids	19

23.	Late Bids	19

24.	Withdrawal, Substitution, and Modification of Bids	19

25.	Bid Opening	20

E. Evaluation and Comparison of Bids	21

26.	Confidentiality	21

27.	Clarification of Bids	21

28.	Deviations, Reservations, and Omissions	22

29.	Determination of Responsiveness	22

30.	Nonconformities, Errors and Omissions	22

31.	Correction of Arithmetical Errors	23

32.	Conversion to Single Currency	23

33.	Margin of  Preference	23

34.	Evaluation of Bids	23

35.	Comparison of Bids	25

36.	Qualification of the Bidder	25

37.	Purchaser’s Right to Accept Any Bid, and to Reject Any or All Bids	25

F. Award of Contract	25

38.	Award Criteria	25

39.	Purchaser’s Right to Vary Quantities at Time of Award	26

40.	Notification of Award	26

41.	Signing of Contract	26

42.	Performance Security	27

| Section I. Instructions to Bidders |  |
| --- | --- |
|  | - General |
| 1. Scope of Bid | - In connection with the Invitation for Bids, specified in the Bid Data Sheet (BDS), the Purchaser, as specified in the BDS, issues these Bidding Documents for the supply of Goods and Related Services incidental thereto as specified in Section VII, Schedule of Requirements. The name, identification and number of lots (contracts) of this International Competitive Bidding (ICB) procurement are specified in the BDS. - Throughout these Bidding Documents: ### the term “in writing” means communicated in written form (e.g. by mail, e-mail, fax, telex) with proof of receipt; ### if the context so requires, “singular” means “plural” and vice versa; and ### “day” means calendar day. |
| 2. Source of Funds | - The Borrower or Recipient (hereinafter called “Borrower”) specified in the BDS has applied for or received financing (hereinafter called “funds”) from the International Bank for Reconstruction and Development or the International Development Association (hereinafter called “the Bank”)in an amount specified in BDS,toward the project named in BDSThe Borrower intends to apply a portion of the funds to eligible payments under the contract for which these Bidding Documents are issued. - Payment by the Bank will be made only at the request of the Borrower and upon approval by the Bank in accordance with the terms and conditions of the Loan (or other financing)Agreement. The Loan (or other financing) Agreement prohibits a withdrawal from the Loan (or other financing) account for the purpose of any payment to persons or entities, or for any import of goods, if such payment or import, to the knowledge of the Bank, is prohibited by decision of the United Nations Security Council taken under Chapter VII of the Charter of the United Nations. No party other than the Borrower shall derive any rights from the Loan (or other financing) Agreement or have any claim to the proceeds of the Loan (or other financing). |
| 3. Corrupt and Fraudulent Practices | 3.1 The Bank requires compliance with its policy in regard to corrupt and fraudulent practices as set forth in Section VI. ### 3.2 In further pursuance of this policy, Bidders shall permit and shall cause its agents (where declared or not), sub-contractors, sub-consultants, service providers or suppliers and to permit the Bank to inspect all accounts, records and other documents relating to the submission of the application, bid submission (in case prequalified), and contract performance (in the case of award), and to have them audited by auditors appointed by the Bank. |
| 4. Eligible Bidders | - A Bidder may be a firm that is a private entity, a government-owned entity—subject to ITB 4.5—or any combination of such entities in the form of a joint venture (JV) under an existing agreement or with the intent to enter into such an agreement supported by a letter of intent. In the case of a joint venture, all members shall be jointly and severally liable for the execution of the Contract in accordance with the Contract terms. The JV shall nominate a Representative who shall have the authority to conduct all business for and on behalf of any and all the members of the JV during the bidding process and, in the event the JV is awarded the Contract, during contract execution.Unless specifiedin the BDS, there is no limit on the number of members in a JV. - A Bidder shall not have a conflict of interest. Any Bidder found to have a conflict of interest shall be disqualified. A Bidder may be considered to have a conflict of interest for the purpose of this bidding process, if the Bidder: ### directly or indirectly controls, is controlled by or is under common control with another Bidder; or ### receives or has received any direct or indirect subsidy from another Bidder; or ### has the same legal representative as another Bidder; or ### has a relationship with another Bidder, directly or through common third parties, that puts it in a position to influence the bid of another Bidder, or influence the decisions of the Purchaser regarding this bidding process; or ### participates in more than one bid in this bidding process. Participation by a Bidder in more than one Bid will result in the disqualification of all Bids in which such Bidder is involved. However, this does not limit the inclusion of the same subcontractor in more than one bid; or ### any of its affiliates participated as a consultant in the preparation of the design or technical specifications of the works that are the subject of the bid; or ### any of its affiliates has been hired (or is proposed to be hired) by the Purchaser or Borrower for the Contract implementation; or ### would be providing goods, works, or non-consulting services resulting from or directly related to consulting services for the preparation or implementation of the project specified in the BDS ITB 2.1 that it provided or were provided by any affiliate that directly or indirectly controls, is controlled by, or is under common control with that firm; or ### has a close business or family relationship with a professional staff of the Borrower (or of the project implementing agency, or of a recipient of a part of the loan) who: (i) are directly or indirectly involved in the preparation of the bidding documents or specifications of the contract, and/or the bid evaluation process of such contract; or (ii) would be involved in the implementation or supervision of such contract unless the conflict stemming from such relationship has been resolved in a manner acceptable to the Bank throughout the procurement process and execution of the contract - ABidder may have the nationality of any country, subject to the restrictions pursuant to ITB 4.7. ABidder shall be deemed to have the nationality of a country if the Bidder is constituted, incorporated or registered in and operates in conformity with the provisions of the laws of that country, as evidenced by its articles of incorporation (or equivalent documents of constitution or association) and its registration documents, as the case may be. This criterion also shall apply to the determination of the nationality of proposed sub-contractors or sub-consultants for any part of the Contract including related Services. - ABidder that has been sanctioned by the Bank in accordance with the above ITB 3.1, including in accordance with the Bank’s Guidelines on Preventing and Combating Corruption in Projects Financed by IBRD Loans and IDA Credits and Grants (“Anti-Corruption Guidelines”), shall be ineligible to be prequalified for, bid for, or be awarded a Bank-financed contract or benefit from a Bank-financed contract, financially or otherwise, during such period of time as the Bank shall have determined. The list of debarred firms and individuals is available at the electronic address specified in the BDS. - Bidders that are Government-owned enterprises or institutions in the Purchaser’s Country may participate only if they can establish that they (i) are legally and financially autonomous (ii) operate under commercial law, and (iii) are not dependent agencies of the Purchaser. To be eligible, a government-owned enterprise or institution shall establish to the Bank’s satisfaction, through all relevant documents, including its Charter and other information the Bank may request, that it: (i) is a legal entity separate from the government (ii) does not currently receive substantial subsidies or budget support; (iii) operates like any commercial enterprise, and, inter alia, is not obliged to pass on its surplus to the government, can acquire rights and liabilities, borrow funds and be liable for repayment of its debts, and can be declared bankrupt; and (iv) is not bidding for a contract to be awarded by the department or agency of the government which under their applicable laws or regulations is the reporting or supervisory authority of the enterprise or has the ability to exercise influence or control over the enterprise or institution. - ABidder shall not be under suspension from bidding by the Purchaser as the result of the operation of a Bid–Securing Declaration. - Firms and individuals may be ineligible if so indicated in Section V and (a) as a matter of law or official regulations, the Borrower’s country prohibits commercial relations with that country, provided that the Bank is satisfied that such exclusion does not preclude effective competition for the supply of goods or the contracting of works or services required; or (b) by an act of compliance with a decision of the United Nations Security Council taken under Chapter VII of the Charter of the United Nations, the Borrower’s country prohibits any import of goods or contracting of works or services from that country, or any payments to any country, person, or entity in that country. - A Bidder shall provide such evidence of eligibility satisfactory to the Purchaser, as the Purchaser shall reasonably request. |
| 5. Eligible Goods and Related Services | - All the Goods and Related Services to be supplied under the Contract and financed by the Bank may have their origin in any country in accordance with Section V, Eligible Countries. - For purposes of this Clause, the term “goods” includes commodities, raw material, machinery, equipment, and industrial plants; and “related services” includes services such as insurance, installation, training, and initial maintenance. - The term “origin” means the country where the goods have been mined, grown, cultivated, produced, manufactured or processed; or, through manufacture, processing, or assembly, another commercially recognized article results that differs substantially in its basic characteristics from its components. |
|  | B. Contents of Bidding Document |
| 6. Sections of Bidding Document | - The Bidding Documents consist of Parts 1, 2, and 3, which include all the Sections indicated below, and should be read in conjunction with any Addenda issued in accordance with ITB 8. PART 1 Bidding Procedures - SectionI. Instructions to Bidders (ITB) - Section II. Bidding Data Sheet (BDS) - Section III. Evaluation and Qualification Criteria - Section IV. Bidding Forms - Section V. Eligible Countries - Section VI. Bank Policy-Corrupt and Fraudulent Practices |
|  | PART 2 Supply Requirements - Section VII. Schedule of Requirements PART 3 Contract - Section VIII. General Conditions of Contract (GCC) - Section IX. Special Conditions of Contract (SCC) - SectionX. Contract Forms |
|  | - The Invitation for Bids issued by the Purchaser is not part of the Bidding Document. - Unless obtained directly from the Purchaser, the Purchaser is not responsible for the completeness of the document, responses to requests for clarification, the Minutes of the pre-Bid meeting (if any), or Addenda to the Bidding Document in accordance with ITB 8. In case of any contradiction, documents obtained directly from the Purchaser shall prevail. - The Bidder is expected to examine all instructions, forms, terms, and specifications in the Bidding Documents and to furnish with its Bid all information or documentation as is required by the Bidding Documents. |
| 7. Clarification of Bidding Documents | - A Bidder requiring any clarification of the Bidding Document shall contact the Purchaser in writing at the Purchaser’s address specified in theBDS. The Purchaser will respond in writing to any request for clarification, provided that such request is received prior to the deadline for submission of bidswithin a period specified in the BDS. The Purchaser shall forward copies of its response to all Bidders who have acquired the Bidding Documents in accordance with ITB 6.3, including a description of the inquiry but without identifying its source. If so specified in the BDS, the Purchaser shall also promptly publish its response at the web page identified in the BDS. Should the clarification result in changes to the essential elements of the Bidding Documents, the Purchaser shall amend the Bidding Documentsfollowing the procedure under ITB 8 and ITB 22.2. |
| 8. Amendment of Bidding Document | - At any time prior to the deadline for submission of bids, the Purchaser may amend the Bidding Documents by issuing addenda. - Any addendum issued shall be part of the Bidding Documents and shall be communicated in writing to all who have obtained the Bidding Documents from the Purchaser in accordance with ITB 6.3.The Purchaser shall also promptly publish the addendum on the Purchaser’s web page in accordance withITB 7.1. - To give prospective Bidders reasonable time in which to take an addendum into account in preparing their bids, the Purchaser may, at its discretion, extend the deadline for the submission of bids, pursuant to ITB 22.2. |
|  | C. Preparation of Bids |
| 9. Cost of Bidding | - The Bidder shall bear all costs associated with the preparation and submission of its bid, and the Purchaser shall not be responsible or liable for those costs, regardless of the conduct or outcome of the bidding process. |
| 10. Language of Bid | - The Bid, as well as all correspondence and documents relating to the bid exchanged by the Bidder and the Purchaser, shall be written in the language specified in the BDS. Supporting documents and printed literature that are part of the Bid may be in another language provided they are accompanied by an accurate translation of the relevant passages into the language specified in theBDS, in which case, for purposes of interpretation of the Bid, such translation shall govern. |
| 11. Documents Comprising the Bid | - The Bid shall comprise the following: ### Letter of Bid in accordance with ITB 12; - completed schedules, in accordance with ITB 12 and 14 ### Bid Security or Bid-Securing Declaration, in accordance with ITB 19.1; ### alternativebids, if permissible, in accordance with ITB 13; ### written confirmation authorizing the signatory of the Bid to commit the Bidder, in accordance with ITB 20.2; ### documentary evidence in accordance with ITB 17 establishing the Bidder’s qualifications to perform the contract if its bid is accepted; ### documentary evidence in accordance with ITB 17 establishing the Bidder’s eligibility to bid; ### documentary evidence in accordance with ITB 16, that the Goods and Related Services to be supplied by the Bidder are of eligible origin; ### documentary evidence in accordance with ITB 16 and 30, that the Goods and Related Services conform to the Bidding Documents; ### any other document required in the BDS. 11.2 In addition to the requirements under ITB 11.1, bids submitted by a JV shall include a copy of the Joint Venture Agreement entered into by all members. Alternatively, a letter of intent to execute a Joint Venture Agreement in the event of a successful bid shall be signed by all members and submitted with the bid, together with a copy of the proposed Agreement. 11.3 The Bidder shall furnish in the Letter of Bid information on commissions and gratuities, if any, paid or to be paid to agents or any other party relating to this Bid. |
| 12. Letter of Bid and Price Schedules | - The Letter of Bid and Price Schedules shall be prepared using the relevant forms furnished in Section IV, Bidding Forms. The forms must be completed without any alterations to the text, and no substitutes shall be accepted except as provided under ITB 20.2. All blank spaces shall be filled in with the information requested. |
| 13. Alternative Bids | - Unless otherwise specified in theBDS, alternative bids shall not be considered. |
| 14. Bid Prices and Discounts | - The prices and discounts quoted by the Bidder in the Letter of Bid and in the Price Schedules shall conform to the requirements specified below. - All lots (contracts) and items must be listed and priced separately in the Price Schedules. - The price to be quoted in the Letter of Bidin accordance with ITB 12.1 shall be the total price of the bid, excluding any discounts offered. - The Bidder shall quote any discounts and indicate the methodology for their application in the Letter of Bid, in accordance with ITB 12.1. - Prices quoted by the Bidder shall be fixed during the Bidder’s performance of the Contract and not subject to variation on any account, unless otherwise specified in the BDSA bid submitted with an adjustable price quotation shall be treated as nonresponsive and shall be rejected, pursuant to ITB 29. However, if in accordance with the BDS, prices quoted by the Bidder shall be subject to adjustment during the performance of the Contract, a bid submitted with a fixed price quotation shall not be rejected, but the price adjustment shall be treated as zero. - If so specified in ITB 1.1, bids are being invited for individual lots (contracts) or for any combination of lots (packages). Unless otherwise specified in the BDS, prices quoted shall correspond to 100 % of the items specified for each lot and to 100% of the quantities specified for each item of a lot. Bidders wishing to offer discounts for the award of more than one Contract shall specify in their bid the price reductions applicable to each package, or alternatively, to individual Contracts within the package. Discounts shall be submitted in accordance with ITB 14.4 provided the bids for all lots (contracts) are opened at the same time. - The terms EXW, CIP, and other similar terms shall be governed by the rules prescribed in the current edition of Incoterms, published by The International Chamber of Commerce, as specified in theBDS. - Prices shall be quoted as specified in each Price Schedule included in Section IV, Bidding Forms. The dis-aggregation of price components is required solely for the purpose of facilitating the comparison of bids by the Purchaser. This shall not in any way limit the Purchaser’s right to contract on any of the terms offered. In quoting prices, the Bidder shall be free to use transportation through carriers registered in any eligible country, in accordance with Section V, Eligible Countries. Similarly, the Bidder may obtain insurance services from any eligible country in accordance with Section V, Eligible Countries. Prices shall be entered in the following manner: ### For Goods manufactured in the Purchaser’s Country: (i) the price of the Goods quoted EXW (ex-works, ex-factory, ex warehouse, ex showroom, or off-the-shelf, as applicable), including all customs duties and sales and other taxes already paid or payable on the components and raw material used in the manufacture or assembly of the Goods; (ii) any Purchaser’s Country sales tax and other taxes which will be payable on the Goods if the contract is awarded to the Bidder; and (iii) the price for inland transportation, insurance, and other local services required to convey the Goods to their final destination (Project Site) specified in theBDS. - For Goods manufactured outside the Purchaser’s Country, to be imported: - the price of the Goods, quoted CIP named place of destination, in the Purchaser’s Country, as specified in theBDS; - the price for inland transportation, insurance, and other local services required to convey the Goods from the named place of destination to their final destination (Project Site) specified in theBDS; - For Goods manufactured outside the Purchaser’s Country, already imported: - the price of the Goods, including the original import value of the Goods; plus any mark-up (or rebate); plus any other related local cost, and custom duties and other import taxes already paid or to be paid on the Goods already imported. - the custom duties and other import taxes already paid (need to be supported with documentary evidence) or to be paid on the Goods already imported; - the price of the Goods, obtained as the difference between (i) and (ii) above; - any Purchaser’s Country sales and other taxes which will be payable on the Goods if the contract is awarded to the Bidder; and - the price for inland transportation, insurance, and other local services required to convey the Goods from the named place of destination to their final destination (Project Site) specified in theBDS. - for Related Services, other than inland transportation and other services required to convey the Goods to their final destination, whenever such Related Services are specified in the Schedule of Requirements: - the price of each item comprising the Related Services (inclusive of any applicable taxes). |
| 15. Currencies of Bid and Payment | - The currency(ies) of the bid and the currency(ies) of payments shall be as specified in the BDS.The Bidder shall quote in the currency of the Purchaser’s Country the portion of the bid price that corresponds to expenditures incurred in the currency of the Purchaser’s country, unless otherwise specified in the BDS. - The Bidder may express the bid price in any currency. If the Bidder wishes to be paid in a combination of amounts in different currencies, it may quote its price accordingly but shall use no more than three foreign currencies in addition to the currency of the Purchaser’s Country. |
| 16. Documents Establishing the Eligibility and Conformity of the Goods and Related Services | - To establish the eligibility of the Goods and Related Services in accordance with ITB 5, Bidders shall complete the country of origin declarations in the Price Schedule Forms, included in Section IV, Bidding Forms. - To establish the conformity of the Goods and Related Services to the Bidding Documents, the Bidder shall furnish as part of its Bid the documentary evidence that the Goods conform to the technical specifications and standards specified in Section VII, Schedule of Requirements. - The documentary evidence may be in the form of literature, drawings or data, and shall consist of a detailed item by item description of the essential technical and performance characteristics of the Goods and Related Services, demonstrating substantial responsiveness of the Goods and Related Services to the technical specification, and if applicable, a statement of deviations and exceptions to the provisions of the Section VII, Schedule of Requirements. - The Bidder shall also furnish a list giving full particulars, including available sources and current prices of spare parts, special tools, etc., necessary for the proper and continuing functioning of the Goods during the period specified in theBDS following commencement of the use of the goods by the Purchaser. - Standards for workmanship, process, material, and equipment, as well as references to brand names or catalogue numbers specified by the Purchaser in the Schedule of Requirements, are intended to be descriptive only and not restrictive. The Bidder may offer other standards of quality, brand names, and/or catalogue numbers, provided that it demonstrates, to the Purchaser’s satisfaction, that the substitutions ensure substantial equivalence or are superior to those specified in the Section VII, Schedule of Requirements. |
| 17. Documents Establishing the Eligibility and Qualifications of the Bidder | - To establish Bidder’s their eligibility in accordance with ITB 4, Bidders shall complete the Letter of Bid, included in Section IV, Bidding Forms. - The documentary evidence of the Bidder’s qualifications to perform the contract if its bid is accepted shall establish to the Purchaser’s satisfaction: - that, if required in theBDS, a Bidder that does not manufacture or produce the Goods it offers to supply shall submit the Manufacturer’s Authorization using the form included in Section IV, Bidding Forms to demonstrate that it has been duly authorized by the manufacturer or producer of the Goods to supply these Goods in the Purchaser’s Country; - that, if required in theBDS,in case of a Bidder not doing business within the Purchaser’s Country, the Bidder is or will be (if awarded the contract) represented by an Agent in the country equipped and able to carry out the Supplier’s maintenance, repair and spare parts-stocking obligations prescribed in the Conditions of Contract and/or Technical Specifications; and - that the Bidder meets each of the qualification criterion specified in Section III, Evaluation and Qualification Criteria. |
| 18. Period of Validity of Bids | - Bids shall remain valid for the period specified in theBDS after the bid submission deadline date prescribed by the Purchaser in accordance with ITB22.1. A bid valid for a shorter period shall be rejected by the Purchaser as nonresponsive. - In exceptional circumstances, prior to the expiration of the bid validity period, the Purchaser may request bidders to extend the period of validity of their bids. The request and the responses shall be made in writing. If a Bid Security is requested in accordance with ITB Clause 19, it shall also be extended for a corresponding period. A Bidder may refuse the request without forfeiting its Bid Security. A Bidder granting the request shall not be required or permitted to modify its bid, except as provided in ITB 18.3. - If the award is delayed by a period exceeding fifty-six (56) days beyond the expiry of the initial bid validity, the Contract price shall be determined as follows: - In the case of fixed price contracts, the Contract price shall be the bid price adjusted by the factor specified in theBDS. - In the case of adjustable price contracts, no adjustment shall be made. - In any case, bid evaluation shall be based on the bid price without taking into consideration the applicable correction from those indicated above. |
| 19. Bid Security | - The Bidder shall furnish as part of its bid, either a Bid-Securing Declaration or a bid security, as specified in theBDS, in original form and, in the case of a bid security.In the amount and currency specified in the BDS. - A Bid Securing Declaration shall use the form included in Section IV, Bidding Forms. - If a bid security is specified pursuant to ITB 19.1, the bid security shall be a demand guarantee in any of the following forms at the Bidder’s option : ### an unconditional guarantee issued by a bank or financial institution (such as an insurance, bonding or surety company); ### an irrevocable letter of credit; ### a cashier’s or certified check; or ### another security specified in the BDS, from a reputable source from an eligible country. If the unconditional guarantee is issued by a financial institution located outside the Purchaser’s Country, the issuing financial institution shall have a correspondent financial institution located in the Purchaser’s Country to make it enforceable. In the case of a bank guarantee, the bid security shall be submitted either using the Bid Security Form included in Section IV, Bidding Forms, or in another substantially similar format approved by the Purchaser prior to bid submission. The bid security shall be valid for twenty-eight (28) days beyond the original validity period of the bid, or beyond any period of extension if requested under ITB 18.2. - If a Bid Security is specified pursuant to ITB 19.1, any bid not accompanied by a substantially responsive Bid Security shall be rejected by the Purchaser as non-responsive. - If a Bid Security is specified pursuant to ITB 19.1, the Bid Security of unsuccessful Bidders shall be returned as promptly as possible upon the successful Bidder’s signing the contract and furnishing the Performance Security pursuant to ITB 42. - The Bid Security of the successful Bidder shall be returned as promptly as possible once the successful Bidder has signed the contract and furnished the required performance security. - The Bid Security may be forfeited or the Bid Securing Declaration executed: ### if a Bidder withdraws its bid during the period of bid validity specified by the Bidder on the Letter of Bid, or any extension thereto provided by the Bidder ; or ### if the successful Bidder fails to: #### sign the Contract in accordance with ITB41; or #### furnish a performance security in accordance with ITB 42. - The bid security or Bid- Securing Declaration of a JV must be in the name of the JV that submits the bid. If the JV has not been legally constituted into a legally enforceable JV at the time of bidding, the bid security or Bid-Securing Declaration shall be in the names of all future members as named in the letter of intent referred to in ITB 4.1 and ITB 11.2. - If a bid security is not required in the BDS, pursuant to ITB 19.1, and - if a Bidder withdraws its bid during the period of bid validity specified by the Bidder on the Letter of Bid, or - if the successful Bidder fails to: sign the Contract in accordance with ITB41; or furnish a performance security in accordance with ITB 42; the Borrower may, if provided for in the BDS, declare the Bidder ineligible to be awarded a contract by the Purchaser for a period of time as stated in the BDS. |
| 20. Format and Signing of Bid | - The Bidder shall prepare one original of the documents comprising the bid as described in ITB 11 and clearly mark it “Original.” Alternative bids, if permitted in accordance with ITB 13, shall be clearly marked “Alternative.” In addition, the Bidder shall submit copies of the bid, in the number specified in the BDS and clearly mark them “Copy.” In the event of any discrepancy between the original and the copies, the original shall prevail. - The original and all copies of the bid shall be typed or written in indelible ink and shall be signed by a person duly authorized to sign on behalf of the Bidder.This authorization shall consist of a written confirmation as specified in the BDS and shall be attached to the bid. The name and position held by each person signing the authorization must be typed or printed below the signature. All pages of the bid where entries or amendments have been made shall be signed or initialed by the person signing the bid. - In case the Bidder is a JV, the Bid shall be signed by an authorized representative of the JV on behalf of the JV, and so as to be legally binding on all the members as evidenced by a power of attorney signed by their legally authorized representatives. - Any inter-lineation, erasures, or overwriting shall be valid only if they are signed or initialed by the person signing the bid. |
|  | D. Submission and Opening of Bids |
| 21. Sealing and Marking of Bids | - The Bidder shall enclose the original and all copies of the bid, including alternative bids, if permitted in accordance with ITB 13, in separate sealed envelopes, duly marking the envelopes as “Original”, “Alternative” and “Copy.” These envelopes containing the original and the copies shall then be enclosed in one single envelope. - The inner and outer envelopes shall: ### bear the name and address of the Bidder; ### be addressed to the Purchaser in accordance with ITB 24.1; ### bear the specific identification of this bidding process indicated in ITB1.1; and ### bear a warning not to open before the time and date for bid opening. - If all envelopes are not sealed and marked as required, the Purchaser will assume no responsibility for the misplacement or premature opening of the bid. |
| 22. Deadline for Submission of Bids | - Bids must be received by the Purchaser at the address and no later than the date and time specifiedin theBDS.When sospecified in the BDS, bidders shall have the option of submitting their bids electronically. Bidders submitting bids electronically shall follow the electronic bid submission procedures specified in the BDS. - The Purchaser may, at its discretion, extend the deadline for the submission of bids by amending the Bidding Documents in accordance with ITB 8, in which case all rights and obligations of the Purchaser and Bidders previously subject to the deadline shall thereafter be subject to the deadline as extended. |
| 23. Late Bids | - The Purchaser shall not consider any bid that arrives after the deadline for submission of bids, in accordance with ITB 22. Any bid received by the Purchaser after the deadline for submission of bids shall be declared late, rejected, and returned unopened to the Bidder. |
| 24. Withdrawal, Substitution, and Modification of Bids | - A Bidder may withdraw, substitute, or modify its Bid after it has been submitted by sending a written notice, duly signed by an authorized representative, and shall include a copy of the authorization (the power of attorney) in accordance with ITB 20.2, (except that withdrawal notices do not require copies ). The corresponding substitution or modification of the bid must accompany the respective written notice. All notices must be: - prepared and submitted in accordance with ITB 20 and 21 (except that withdrawal notices do not require copies), and in addition, the respective envelopes shall be clearly marked “Withdrawal,” “Substitution,” or “Modification;” and - received by the Purchaser prior to the deadline prescribed for submission of bids, in accordance with ITB 22. - Bids requested to be withdrawn in accordance with ITB 24.1 shall be returned unopened to the Bidders. - No bid may be withdrawn, substituted, or modified in the interval between the deadline for submission of bids and the expiration of the period of bid validity specified by the Bidder on the Letter of Bid or any extension thereof. |
| 25. Bid Opening | - Except as in the cases specified in ITB 23 and 24, the Purchaser shall publicly open and read out in accordance with ITB25.3 all bids received by the deadline at the date, time and placespecified in theBDSin the presence of Bidders’designated representatives and anyone who choose to attend. Any specific electronic bid opening procedures required if electronic bidding is permitted in accordance with ITB 22.1, shall be as specified in theBDS. - First, envelopes marked “Withdrawal” shall be opened and read out and the envelope with the corresponding bid shall not be opened, but returned to the Bidder. If the withdrawal envelope does not contain a copy of the “power of attorney” confirming the signature as a person duly authorized to sign on behalf of the Bidder, the corresponding bid will be opened. No bid withdrawal shall be permitted unless the corresponding withdrawal notice contains a valid authorization to request the withdrawal and is read out at bid opening. Next, envelopes marked “Substitution” shall be opened and read out and exchanged with the corresponding Bid being substituted, and the substituted Bid shall not be opened, but returned to the Bidder. No Bid substitution shall be permitted unless the corresponding substitution notice contains a valid authorization to request the substitution and is read out at bid opening. Envelopes marked “Modification” shall be opened and read out with the corresponding Bid. No Bid modification shall be permitted unless the corresponding modification notice contains a valid authorization to request the modification and is read out at Bid opening. Only bids that are opened and read out at Bid opening shall be considered further. - All other envelopes shall be opened one at a time, reading out: the name of the Bidder and whether there is a modification; the total Bid Prices, per lot (contract) if applicable, including any discounts and alternative bids; the presence or absence of a Bid Security, if required; and any other details as the Purchaser may consider appropriate. Only discounts and alternative bids read out at Bid opening shall be considered for evaluation. The Letter of Bid and the Price Schedules are to be initialed by representatives of the Purchaser attending bid opening in the manner specified in theBDS.The Purchasershall neither discuss the merits of any bid nor reject any bid (except for late bids, in accordance with ITB 25.1). - The Purchaser shall prepare a record of the bid opening that shall include, as a minimum: the name of the Bidder and whether there is a withdrawal, substitution, or modification; the Bid Price, per lot (contract) if applicable, including any discounts, and alternative bids; and the presence or absence of a Bid Security, if one was required. The Bidders’ representatives who are present shall be requested to sign the record. The omission of a Bidder’s signature on the record shall not invalidate the contents and effect of the record. A copy of the record shall be distributed to all Bidders. |
|  | E. Evaluation and Comparison of Bids |
| 26. Confidentiality | - Information relating to the evaluation of bids andrecommendation of contract award, shall not be disclosed to bidders or any other persons not officially concerned with the bidding process until information on Contract Award is communication to all Bidders in accordance with ITB 40. - Any effort by a Bidder to influence the Purchaser in the evaluation or contract award decisions may result in the rejection of its Bid. - Notwithstanding ITB 26.2, from the time of bid opening to the time of Contract Award, if any Bidder wishes to contact the Purchaser on any matter related to the bidding process, it should do so in writing. |
| 27. Clarification of Bids | - To assist in the examination, evaluation, comparison of the bids, and qualification of the Bidders, the Purchaser may, at its discretion, ask any Bidder for a clarificationof its Bid. Any clarification submitted by a Bidder in respect to its Bid and that is not in response to a request by the Purchaser shall not be considered. The Purchaser’s request for clarification and the response shall be in writing. No change, including any voluntary increase or decrease, in the prices or substance of the Bid shall be sought, offered, or permitted, except to confirm the correction of arithmetic errors discovered by the Purchaser in the Evaluation of the bids, in accordance with ITB 31. - If a Bidder does not provide clarifications of its bid by the date and time set in the Purchaser’s request for clarification, its bid may be rejected. |
| 28. Deviations, Reservations, and Omissions | - During the evaluation of bids, the following definitions apply: - “Deviation” is a departure from the requirements specified in the Bidding Documents; - “Reservation” is the setting of limiting conditions or withholding from complete acceptance of the requirements specified in the Bidding Documents; and - “Omission” is the failure to submit part or all of the information or documentation required in the Bidding Documents |
| 29. Determination of Responsiveness | - The Purchaser’s determination of a bid’s responsiveness is to be based on the contents of the bid itself, as defined in ITB 11. - A substantially responsive Bid is one that meets the requirements of the Bidding Documents without material deviation, reservation, or omission. A material deviation, reservation, or omission is one that: ### if accepted, would ### affect in any substantial way the scope, quality, or performance of the Goods and Related Services specified in the Contract; or ### limit in any substantial way, inconsistent with the Bidding Documents, the Purchaser’s rights or the Bidder’s obligations under the Contract; or ### if rectified, would unfairly affect the competitive position of other bidders presenting substantially responsive bids. - The Purchaser shall examine the technical aspects of the bid submitted in accordance with ITB 16 and ITB 17, in particular, to confirm that all requirements of Section VII, Schedule of Requirements have been met without any material deviation or reservation, or omission. - If a bid is not substantially responsive to the requirements of Bidding Documents, it shall be rejected by the Purchaser and may not subsequently be made responsive by correction of the material deviation, reservation, or omission. |
| 30. Nonconformities, Errors and Omissions | - Provided that a Bid is substantially responsive, the Purchaser may waive any nonconformities in the Bid. - Provided that a bid is substantially responsive, the Purchaser may request that the Bidder submit the necessary information or documentation, within a reasonable period of time, to rectify nonmaterial nonconformities or omissions in the bid related to documentation requirements. Such omission shall not be related to any aspect of the price of the Bid. Failure of the Bidder to comply with the request may result in the rejection of its Bid. - Provided that a bid is substantially responsive, the Purchaser shall rectify quantifiable nonmaterial nonconformities related to the Bid Price. To this effect, the Bid Price shall be adjusted, for comparison purposes only, to reflect the price of a missing or non-conforming item or component. |
| 31. Correction of Arithmetical Errors | - Provided that the Bid is substantially responsive, the Purchaser shall correct arithmetical errors on the following basis: ### if there is a discrepancy between the unit price and the line item total that is obtained by multiplying the unit price by the quantity, the unit price shall prevail and the line item total shall be corrected, unless in the opinion of the Purchaser there is an obvious misplacement of the decimal point in the unit price, in which case the line item total as quoted shall govern and the unit price shall be corrected; ### if there is an error in a total corresponding to the addition or subtraction of subtotals, the subtotals shall prevail and the total shall be corrected; and ### if there is a discrepancy between words and figures, the amount in words shall prevail, unless the amount expressed in words is related to an arithmetic error, in which case the amount in figures shall prevail subject to (a) and (b) above. - Bidders shall be requested to accept correction of arithmetical errors. Failure to accept the correction in accordance with ITB 31.1, shall result in the rejection of the Bid. |
| 32. Conversion to Single Currency | - For evaluation and comparison purposes, the currency(ies) of the Bid shall be converted in a single currency as specified in theBDS. |
| 33. Margin of Preference | - Unless otherwise specified in theBDS, a margin of preference shall not apply. |
| 34. Evaluation of Bids | - The Purchaser shall use the criteria and methodologies listedin this Clause. No other evaluation criteria or methodologies shall be permitted. - To evaluate a Bid, the Purchaser shall consider the following: ### evaluation will be done for Items or Lots (contracts), as specified in theBDS; and the Bid Price as quoted in accordance with clause 14; ### price adjustment for correction of arithmetic errors in accordance with ITB 31.1; ### price adjustment due to discounts offered in accordance with ITB 14.3; ### converting the amount resulting from applying (a) to (c) above, if relevant, to a single currency in accordance with ITB 32; ### price adjustment due to quantifiable nonmaterial nonconformities in accordance with ITB 30.3; ### the additional evaluation factors are specified in Section III, Evaluation and Qualification Criteria; - The estimated effect of the price adjustment provisions of the Conditions of Contract, applied over the period of execution of the Contract, shall not be taken into account in bid evaluation. - If these Bidding Documents allows Bidders to quote separate prices for different lots (contracts), the methodology to determine the lowest evaluated price of the lot (contract) combinations, including any discounts offered in the Letter of Bid Form, is specified in Section III, Evaluation and Qualification Criteria - The Purchaser’s evaluation of a bid will exclude and not take into account: ### in the case of Goods manufactured in the Purchaser’s Country, sales and other similar taxes, which will be payable on the goods if a contract is awarded to the Bidder; ### in the case of Goods manufactured outside the Purchaser’s Country, already imported or to be imported, customs duties and other import taxes levied on the imported Good, sales and other similar taxes, which will be payable on the Goods if the contract is awarded to the Bidder; ### any allowance for price adjustment during the period of execution of the contract, if provided in the bid. - The Purchaser’s evaluation of a bid may require the consideration of other factors, in addition to the Bid Price quoted in accordance with ITB 14. These factors may be related to the characteristics, performance, and terms and conditions of purchase of the Goods and Related Services. The effect of the factors selected, if any, shall be expressed in monetary terms to facilitate comparison of bids, unless otherwise specified in the BDS from amongst those set out in Section III, Evaluation and Qualification Criteria. The criteria and methodologies to be used shall be as specified in ITB 34.2 (f). |
| 35. Comparison of Bids | - The Purchaser shall compare the evaluated prices of all substantially responsive bids established in accordance with ITB 34.2to determine the lowestevaluated bid.The comparison shall be on the basis of CIP (place of final destination) prices for imported goods and EXW prices, plus cost of inland transportation and insurance to place of destination, for goods manufactured within the Borrower’s country, together with prices for any required installation, training, commissioning and other services.The evaluation of prices shall not take into accountcustom duties and other taxes levied on imported goods quoted CIP and sales and similar taxes levied in connection with the sale or delivery of goods. |
| 36. Qualification of the Bidder | - The Purchaser shall determine to its satisfaction whether the Bidder that is selected as having submitted the lowest evaluated and substantially responsive bid meets the qualifying criteria specified in Section III, Evaluation and Qualification Criteria. - The determination shall be based upon an examination of the documentary evidence of the Bidder’s qualifications submitted by the Bidder, pursuant to ITB 17. - An affirmative determination shall be a prerequisite for award of the Contract to the Bidder. A negative determination shall result in disqualification of the bid, in which event the Purchaser shall proceed to the next lowest evaluated bid to make a similar determination of that Bidder’s qualifications to perform satisfactorily. |
| 37. Purchaser’s Right to Accept Any Bid, and to Reject Any or All Bids | - The Purchaser reserves the right to accept or reject any bid, and to annul the bidding process and reject all bids at any time prior to contract award, without thereby incurring any liability to Bidders.In case of annulment, all bids submitted and specifically, bid securities, shall be promptly returned to the Bidders. |
|  | F. Award of Contract |
| 38. Award Criteria | - Subject to ITB 37.1, the Purchaser shall award the Contract to the Bidder whose bid has been determined to be the lowest evaluated bid and is substantially responsive to the Bidding Documents, provided further that the Bidder is determined to be qualified to perform the Contract satisfactorily. |
| 39. Purchaser’s Right to Vary Quantities at Time of Award | - At the time the Contract is awarded, the Purchaser reserves the right to increase or decrease the quantity of Goods and Related Services originally specified in Section VII, Schedule of Requirements, provided this does not exceed the percentages specified in the BDS, and without any change in the unit prices or other terms and conditions of the bid and the Bidding Documents. |
| 40. Notification of Award | - Prior to the expiration of the period of bid validity, the Purchaser shall notify the successful Bidder, in writing, that its Bid has been accepted. The notification letter (hereinafter and in the Conditions of Contract and Contract Forms called the “Letter of Acceptance”) shall specify the sum that the Purchaser will pay the Supplierin consideration of the supply of Goods (hereinafter and in the Conditions of Contract and Contract Forms called “the Contract Price”). At the same time, the Purchaser shall also notify all other Bidders of the results of the bidding and shall publish in UNDB online the results identifying the bid and lot (contract) numbers and the following information: (i) name of each Bidder who submitted a Bid; (ii) bid prices as read out at Bid Opening; (iii) name and evaluated prices of each Bid that was evaluated; (iv) name of bidders whose bids were rejected and the reasons for their rejection; and (v) name of the successful Bidder, and the Price it offered, as well as the duration and summary scope of the contract awarded. - Until a formal Contract is prepared and executed, the notification of award shall constitute a binding Contract. - The Purchaser shall promptly respond in writing to any unsuccessful Bidder who, after notification of award in accordance with ITB 40.1, requests in writing the grounds on which its bid was not selected. |
| 41. Signing of Contract | - Promptly after notification, the Purchaser shall send the successful Bidder the Contract Agreement. - Within twenty-eight (28) days of receipt of the Contract Agreement, the successful Bidder shall sign, date, and return it to the Purchaser. - Notwithstanding ITB 41.2 above, in case signing of the Contract Agreement is prevented by any export restrictions attributable to the Purchaser, to the country of the Purchaser, or to the use of the products/goods, systems or services to be supplied, where such export restrictions arise from trade regulations from a country supplying those products/goods, systems or services, the Bidder shall not be bound by its bid, always provided however, that the Bidder can demonstrate to the satisfaction of the Purchaserand of the Bank that signing of the Contact Agreement has not been prevented by any lack of diligence on the part of the Bidder in completing any formalities, including applying for permits, authorizations and licenses necessary for the export of the products/goods, systems or services under the terms of the Contract. |
| 42. Performance Security | - Within twenty eight (28) days of the receipt of notification of award from the Purchaser, the successful Bidder, if required, shall furnish the Performance Security in accordance with the GCC, subject to ITB 34.5, using for that purpose the Performance Security Form included in Section X, Contract Forms, or another Form acceptable to the Purchaser. If the Performance Security furnished by the successful Bidder is in the form of a bond, it shall be issued by a bonding or insurance company that has been determined by the successful Bidder to be acceptable to the Purchaser. A foreign institution providing a bond shall have a correspondent financial institution located in the Purchaser’s Country. - Failure of the successful Bidder to submit the above-mentioned Performance Security or sign the Contract shall constitute sufficient grounds for the annulment of the award and forfeiture of the Bid Security. In that event the Purchaser may award the Contract to the next lowest evaluated Bidder, whose bid is substantially responsive and is determined by the Purchaser to be qualified to perform the Contract satisfactorily. |

| Section II. Bid Data Sheet (BDS) The following specific data for the goods to be procured shall complement, supplement, or amend the provisions in the Instructions to Bidders (ITB). Whenever there is a conflict, the provisions herein shall prevail over those in ITB. |  |
| --- | --- |
| ITB Clause Reference | A. General |
| ITB 1.1 | The reference number of the Invitation for Bids is: CR4/ICB/B-G/013-16 |
| ITB 1.1 | The Purchaser is: “Health projects implementation unit” SA |
| ITB 1.1 | The name of the ICB is: Medical, diagnostic equipment, medical supplies and furniture for “Vanadzor MC” CJSC, comprising of 4 lots The identification number of the ICB is: CR4/ICB/B-G/013-16 The number and identification of lots (contracts)comprising this ICB is: Lot 1- Supply and Installation of Diagnostic and Imaging equipment for "Vanadzor MC” CJSC, No: CR4/ICB/B-G/013-16/1 Lot 2- Supply and Installation of Medical Laboratory equipment for "Vanadzor MC” CJSC, No: CR4/ICB/B-G/013-16/2 Lot 3 - Supply and Installation of Surgery and Hospital equipment for "Vanadzor MC” CJSC, No: CR4/ICB/B-G/013-16/3 Lot 4 -Supply and Installation of Medical Furniture for "Vanadzor MC” CJSC, No: CR4/ICB/B-G/013-16/4 |
| ITB 2.1 | The Borrower is: Republic of Armenia, “Health projects implementation unit” SA, Ministry of Health. |
| ITB 2.1 | Loan or Financing Agreement amount: USD 35,000,000 |
| ITB 2.1 | The name of the Project is: Disease prevention and control project Credit No. 5222-AM |
| ITB 4.1 | Maximum number of members in the JV shall be: 2 (two) |
| - IITB 4.4 | A list of debarred firms and individuals is available on the Bank’s external website: http://www.worldbank.org/debarr. |
|  | B. Contents of Bidding Documents |
| ITB 7.1 | For Clarification of bid purposes only, the Purchaser’s address is: Attention: Hovakim Podosyan, procurement specialist of the Health PIU SA Address: Komitas 49/4 avenue Floor/ Room number: 5th floor, procurement department City: Yerevan ZIP Code: 0051 Country: Republic of Armenia Telephone: +374 10-297-536, +374 10-297-537, Electronic mail address: info@healthpiu.am; procurement@healthpiu.am Requests for clarification should be received by the Employer no later than: [7 days] before Bid opening. |
| ITB 7.1 | Web page: www.procurement.am (www.gnumner.am) |
|  | C. Preparation of Bids |
| ITB 10.1 | The language of the bid is: “English”. All correspondence exchange shall be in “English” language. Language for translation of supporting documents and printed literature is Armenian language. |
| ITB 11.1 (i) | The Bidder shall furnish a detailed comprising table of suggested technical specifications of the goods (electronic version (USB or CD, in excel format). |
| ITB 11.1 (j) | The Bidder shall submit the following additional documents in its bid: - Statement from Manufacturers or Bidder that the goods (offered models) supplied in the frames of the contract should be manufactured not earlier than 12 month prior the bid opening data. - For Lots 1, 2 and 3 applicable: Documentary evidence (copy of certificates) that the goods are qualified with ISO13485 (or equivalent). The goods must be also certified with CE Mark according to 93/42/EEC (or equivalent), according to the technical specifications. |
| ITB 13.1 | Alternative Bids “shall not be” considered. |
| ITB 14.5 | The prices quoted by the Bidder “shall not” be subject to adjustment during the performance of the Contract. |
| ITB 14.6 | Prices quoted for each lot (contract) shall correspond at least to [100] percent of the items specified for each lot (contract). Prices quoted for each item of a lot shall correspond at least to [100] percent of the quantities specified for this item of a lot. |
| ITB 14.7 | The Incoterms edition is: INCOTERMS 2010. |
| ITB 14.8 (b) (i) and (c) (v) | Place of Destination: CIP Yerevan |
| ITB 14.8 (a) (iii);(b)(ii) and (c)(v) | “Final destination (Project Site)”: Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |
| ITB 15.1 | The prices shall be quoted by the bidder in: “The bidder may express the bid price in any currency. If the bidder wishes to express the bid as a sum of amounts in different foreign currencies, they may do so, provided the price expressed no more than three foreign currencies”. The Bidder “is not” required to quote in the currency of the Purchaser’s Country the portion of the bid price that corresponds to expenditures incurred in that currency. The local bidders are reminded that the requirements of “The Low of the currency regulation and currency control”# 135-N, adopted by the National Assembly of RA on 24.11.2004, should be respected. |
| ITB 16.4 | Period of time the Goods are expected to be functioning (for the purpose of spare parts): Four (4) years. |
| ITB 17.2 (a) | Manufacturer’s authorization is: The original MAFs are “required” for all items, under Lots 1; 2; 3. Under Lot 4 original MAFs are required only for items 9-10; 12; 14-15; 17-18; 20-24; 29-30; 32-33; 37; 41-42. |
| ITB 17.2 (b) | After sales service is: “required” |
| ITB 18.1 | The bid validity period shall be 90 (ninety) days. |
| ITB 18.3 (a) | The bid price shall be adjusted by the following factor(s): 1.02 |
| ITB 19.1 | A Bid Security “shall be” required. A Bid-Securing Declaration “shall not be” required. The amount and currency of the bid security shall be as follow: Lot 1: USD 26,500 Lot 2: USD 1,600 Lot 3: USD 26,500 Lot 4: USD 6,600 Note: Bid Security is required for each lot as per amounts indicated against each lot. Bidders have the option of submitting one Bid Security for all lots (for the combined total amount of all lots) for which Bids have been submitted, however if the amount of Bid Security is less than the total required amount, the Purchaser will determine for which lot or lots the Bid Security amount shall be applied.] |
| ITB 19.3 (b) (c) | Not applicable |
| ITB 19.3 (d) | Other types of acceptable securities: None. |
| ITB 19.9 | If the Bidder incurs any of the actions prescribed in subparagraphs (a) or (b) of this provision, the Borrower will declare the Bidder ineligible to be awarded contracts by the Purchaser for a period of_______Not applicable. |
| ITB 20.1 | In addition to the original of the bid, the number of copies is: 2 (two) |
| ITB 20.2 | The written confirmation of authorization to sign on behalf of the Bidder shall consist of: Authorization shall consist of a written confirmation from the company’s director or president and shall be attached to the bid. |
|  | D. Submission and Opening of Bids |
| ITB 22.1 | For bid submission purposes only, the Purchaser’s address is: Attention: Nelson Zuloyan, Acting Director of the Health PIU SA Street Address: Komitas 49/4 ave. Floor/ Room number: 5th floor, procurement department City: Yerevan ZIP/Postal Code: 0051 Country: Republic of Armenia The deadline for bid submission is: Date: 01 March, 2017 Time: 10:00 a.m.(time zone: UTC/GMT +04:00 hours) Bidders “shall not” have the option of submitting their bids electronically. If bidders have the option of submitting their bids electronically, the electronic bidding submission procedures shall be: Not applicable |
| ITB 25.1 | The bid opening shall take place at: Street Address: Komitas 49/4 ave., 5th floor, Meeting room Floor/ Room number: 5th floor, Meeting room City: Yerevan Country: Republic of Armenia Date: 01 March, 2017 Time: 10:00 a.m.(time zone: UTC/GMT +04:00 hours) If bidders have the option of submitting their bids electronically, the electronic bidding opening procedures shall be: Not applicable |
| ITB 25.3 | The Letter of Bid and Price Schedules shall be initialed by one representatives of the Purchaser conducting Bid opening. Each Bid will be initialed by one member of evaluation committee and will be numbered, any modification to the unit or total price will be initialed by the committee member. |
| E. Evaluation and Comparison of Bids |  |
| ITB 32.1 | The currency that shall be used for bid evaluation and comparison purposes to convert all bid prices expressed in various currencies into a single currency is: Armenian drams The source of exchange rate shall be: Central Bank of the Republic of Armenia. www.cba.am The date for the exchange rate shall be: Bid opening day. |
| ITB 33.1 | A margin of domestic preference “shall not” apply. If a margin of preference applies, the application methodology shall be defined in Section III – Evaluation and Qualification Criteria. Not applicable |
| ITB 34.2(a) | Evaluation will be done for Lots (contracts) Note: Bids will be evaluated lot by lot. If a Price Schedule shows items listed but not priced, their prices shall be assumed to be included in the prices of other items. An item not listed in the Price Schedule shall be assumed to be not included in the bid, and provided that the bid is substantially responsive, the average price of the item quoted by substantially responsive bidders will be added to the bid price and the equivalent total cost of the bid so determined will be used for price comparison. |
| ITB 34.6 | The adjustments shall be determined using the following criteria, from amongst those set out in Section III, Evaluation and Qualification Criteria: - Deviation in Delivery schedule: No. - Deviation in payment schedule: No. - the cost of major replacement components, mandatory spare parts, and service: No. - the availability in the Purchaser’s Country of spare parts and after-sales services for the equipment offered in the bid No. - the projected operating and maintenance costs during the life of the equipment No. - the performance and productivity of the equipment offered; No. |
|  | F. Award of Contract |
| ITB 39.1 | The maximum percentage by which quantities may be increased is: 20 The maximum percentage by which quantities may be decreased is: 20 |

Section III.  Evaluation and Qualification Criteria

This Section contains all the criteria that the Purchaser shall use to evaluate a bid and qualify the Bidders. in accordance with ITB 34 and ITB 36, no other factors, methods or criteria shall be used.

Contents

1. Margin of Preference (ITB 33)	38

2. Evaluation(ITB 34)	38

3. Qualification(ITB 36)	38

1. Margin of Preference (ITB 33)

If the Bidding Data Sheet so specifies, the Purchaser will grant a margin of preference to goods manufactured in the Purchaser’s country for the purpose of bid comparison, in accordance with the procedures outlined in subsequent paragraphs. Not applicable

2. Evaluation(ITB 34)

2.1. Evaluation Criteria (ITB 34.6)

The Purchaser’s evaluation of a bid may take into account, in addition to the Bid Price quoted in accordance with ITB Clause 14.8, one or more of the following factors as specified in ITB34.2 (f) and in BDS referring to ITB34.6, using the following criteria and methodologies. Not applicable

2.2. Multiple Contracts (ITB 34.4)

The Purchaser shall award multiple contracts to the Bidder that offers the lowest evaluated combination of bids (one contract per bid) and meets the post-qualification criteria (this Section III, Sub-Section ITB 36.1 Post-Qualification Requirements)

The Purchaser shall:

(a)	evaluate only lots or contracts that include at least the percentages of items per lot and quantity per item as specified in ITB 14.8

(b)	take into account:

      - the lowest-evaluated bid for each lot and

(ii)	the price reduction per lot and the methodology for its  application as offered by the Bidder in its bid”

2.3. Alternative Bids (ITB 13.1)

An alternative if permitted under ITB 13.1, will be evaluated as follows: Not applicable

3. Qualification (ITB 36)

3.1 Post qualification Requirements (ITB 36.1)

After determining the lowest-evaluated bid in accordance with ITB 35.1, the Purchaser shall carry out the post qualification of the Bidder in accordance with ITB 36, using only the requirements specified.  Requirements not included in the text below shall not be used in the evaluation of the Bidder’s qualifications.

    - If bidder is Manufacturer:

- Financial Capability: The Bidder shall furnish documentary evidence that it meets the following financial requirement(s): Having financial resources or access to credit line equal to 80% of proposed bid price.

- Experience and Technical Capacity: The Bidder shall furnish documentary evidence to demonstrate that it meets the following experience requirement(s): For Lot 1; Lot 2 and Lot 3 -At least 5 years experience in manufacturing of Medical equipment. For Lot 4 -At least 5 years experience in manufacturing of Medical furniture.

- Documentary Evidence: The Bidder shall furnish documentary evidence to demonstrate that the Goods it offers meet the following usage requirement:

- For Lot 1-Reference letter from beneficers that offered goods have been in succesful use at least for 6 months (except item 7).

- For Lot 2-Reference letter from beneficers that offered goods have been in succesful use at least for 6 months, required only for items 1-4; 14-15.

- For Lot 3-Reference letter from beneficers that offered goods have been in succesful use at least for 6 months, required only for items 3-5; 7; 11; 20; 26; 32; 34-36.

- For Lot 4-Reference letter from beneficers that offered goods have been in succesful use at least for 6 months, required only for items 9-10; 14-15.

The reference letters could be issued on the basis of beneficiaries’ cooperation’s with Bidder/s or manufacturer/s.

- Similar experience: The Bidder shall furnish documentary evidence (hard copy of similar contract) to demonstrate that:

For Lot 1 -Has successfully completed at least 1 contract of supply Medical Diagnostic equipment in the past 5 years (2012-2016). Contract price should be not less than USD 700,000 (tax inclusive) or equal in other currency.The submitted contract should be accompanied with final acceptance act/final payment evidence/references from beneficiaries on successful completion of contracts.

For Lot 2 -Has successfully completed at least 1 contract of Medical Laboratory equipment in the past 5 years (2012-2016). Contract price should be not less than USD 66,000 (tax inclusive) or equal in other currency.The submitted contract should be accompanied with final acceptance act/final payment evidence/references from beneficiaries on successful completion of contracts.

For Lot 3 -Has successfully completed at least 1 contract of supply of Surgery equipment in the past 5 years (2012-2016). Contract price should be not less than USD 700,000 (tax inclusive) or equal in other currency. The submitted contract should be accompanied with final acceptance act/final payment evidence/references from beneficiaries on successful completion of contracts.

For Lot 4 -Has successfully completed at least 1 contract of Medical furniture in the past 5 years (2012-2016). Contract price should be not less than USD 265,000 (tax inclusive) or equal in other currency.The submitted contract should be accompanied with final acceptance act/final payment evidence/references from beneficiaries on successful completion of contracts.

    - If Bidder is not manufacturer:

If a Bidder is not a manufacturer, but is offering the Goods on behalf of the Manufacturer under Manufacturer's Authorization Form (Section IV, Bidding Forms), the Manufacturer shall demonstrate the above qualifications (ii) and the Bidder shall demonstrate that it has above qualifications (i); (iii) and (iv).

The Bidder or the agent in case of submitting lot 1 or/and lot 2 or/and lot 3 shall have at least 3 years experience in providing of installation, operational trainings and technical and professional services of Medical equipment

Note: If the Bidder submits Bids for more than one Lot, he should comply to aggregate requirements (financial capacity and similar experience).

Section IV. Bidding Forms

Table of Forms

Letter of Bid	42

Bidder Information Form	45

Bidder’s JV Members Information Form	46

Price Schedule: Goods Manufactured Outside the Purchaser’s Country, to be Imported	48

Price Schedule: Goods Manufactured Outside the Purchaser’s Country, already imported*	49

Price Schedule: Goods Manufactured in the Purchaser’s Country	50

Price and Completion Schedule - Related Services	51

Form of Bid Security	52

Form of Bid Security (Bid Bond)	54

Form of Bid-Securing Declaration	Error! Bookmark not defined.

Manufacturer’s Authorization	56

Letter of Bid

The Bidder must prepare the Letter of Bid on stationery with its letterhead clearly showing the Bidder’s complete name and address. Note: All italicized text is for use in preparing these form and shall be deleted from the final products.

Date: [insert date (as day, month and year) of Bid Submission]

ICB No.: [insert number of bidding process]

Invitation for Bid No.: [insert identification]

Alternative No.:[insert identification No if this is a Bid for an alternative]

To:  [insert complete name of Purchaser]

- We have examined and have no reservations to the Bidding Documents, including Addenda issued in accordance with Instructions to Bidders (ITB 8)	;

- We meet the eligibility requirements and have no conflict of interest in accordance with ITB 4;

- We have not been suspended nor declared ineligible by the Purchaser  based on execution of a Bid Securing Declaration in the Purchaser’s country in accordance with ITB 4.6

- We offer to supply in conformity with the Bidding Documents and in accordance with the Delivery Schedules specified in the Schedule of Requirements the following Goods:[insert a brief description of the Goods and Related Services];

- The total price of our Bid, excluding any discounts offered in item (f) below is:

In case of only one lot, total price of the Bid [insert the total price of the bid in words and figures, indicating the various amounts and the respective currencies];

In case of multiple lots, total price of each lot [insert the total price of each lot in words and figures, indicating the various amounts and the respective currencies];

In case of multiple lots, total price of all lots (sum of all lots) [insert the total price of all lots in words and figures, indicating the various amounts and the respective currencies];

- The discounts offered and the methodology for their application are:

(i) The discounts offered are: [Specify in detail each discount offered.]

(ii) The exact method of calculations to determine the net price after application of discounts is shown below:[Specify in detail the method that shall be used to apply the discounts];

- Our bid shall be valid for a period of [specify the number of calendar days]  days from the date fixed for the bid submission deadline in accordance with the Bidding Documents, and it shall remain binding upon us and may be accepted at any time before the expiration of that period;

- If our bid is accepted, we commit to obtain a performance security in accordance with the Bidding Documents;

- We are not participating, as a Bidder or as a subcontractor, in more than one bid in this bidding process in accordance with ITB 4.2(e), other than alternative bids submitted in accordance with ITB 13;

- We, along with any of our subcontractors, suppliers, consultants, manufacturers, or service providers for any part of the contract, are not subject to, and not controlled by any entity or individual that is subject to, a temporary suspension or a debarment imposed by a member of the World Bank Group or a debarment imposed by the World Bank Group in accordance with the Agreement for Mutual Enforcement of Debarment Decisions between the World Bank and other development banks. Further, we are not ineligible under the Employer’s country laws or official regulations or pursuant to a decision of the United Nations Security Council;

- We are not a government owned entity/ We are a government owned entity but meet the requirements of ITB 4.5;

- We have paid, or will pay the following commissions, gratuities, or fees with respect to the bidding process or execution of the Contract: [insert complete name of each Recipient, its full address, the reason for which each commission or gratuity  was paid and the amount and currency of each such commission or gratuity]

| Name of Recipient | Address | Reason | Amount |
| --- | --- | --- | --- |

(If none has been paid or is to be paid, indicate “none.”)

- We understand that this bid, together with your written acceptance thereof included in your notification of award, shall constitute a binding contract between us, until a formal contract is prepared and executed; and

- We understand that you are not bound to accept the lowest evaluated bid or any other bid that you may receive.

- We hereby certify that we have taken steps to ensure that no person acting for us or on our behalf will engage in any type of fraud and corruption

Name of the Bidder*	[insert complete name of Bidder]

Name of the person duly authorized to sign the Bid on behalf of the Bidder** [insert complete name of person duly authorized to sign the Bid]

Title of the person signing the Bid [insert complete title of the person signing the Bid]

Signature of the person named above	 [insert signature of person whose name and capacity are shown above]

Date signed _[insert date of signing] day of [insert month], [insert year]

*: In the case of the Bid submitted by joint venture specify the name of the Joint Venture as Bidder

**: Person signing the Bid shall have the power of attorney given by the Bidder to be attached with the Bid Schedules.

Bidder Information Form

[The Bidder shall fill in this Form in accordance with the instructions indicated below. No alterations to its format shall be permitted and no substitutions shall be accepted.]

Date: [insert date (as day, month and year) of Bid Submission]

ICB No.: [insert number of bidding process]

Alternative No.: [insert identification No if this is a Bid for an alternative]

Page ________ of_ ______ pages

1. Bidder’s Name [insert Bidder’s legal name]

2. In case of JV, legal name of each member : [insert legal name of each member in JV]

3. Bidder’s actual or intended country of registration: [insert actual or intended country of registration]

4. Bidder’s year of registration: [insert Bidder’s year of registration]

5. Bidder’s Address in country of registration: [insert Bidder’s legal address in country of registration]

6. Bidder’s Authorized Representative Information Name: [insert Authorized Representative’s name] Address: [insert Authorized Representative’s Address] Telephone/Fax numbers: [insert Authorized Representative’s telephone/fax numbers] Email Address: [insert Authorized Representative’s email address]

7. Attached are copies of original documents of[check the box(es) of the attached original documents] Articles of Incorporation (or equivalent documents of constitution or association), and/or documents of registration of the legal entity named above, in accordance with ITB 4.3. In case of JV, letter of intent to form JV or JV agreement, in accordance with ITB 4.1. In case of Government-owned enterprise or institution, in accordance with ITB 4.5 documents establishing: - Legal and financial autonomy - Operation under commercial law - Establishing that the Bidder is not dependent agency of the Purchaser 2. Included are the organizational chart, a list of Board of Directors, and the beneficial ownership.

Bidder’s JV Members Information Form

[The Bidder shall fill in this Form in accordance with the instructions indicated below. The following table shall be filled in for the Bidder and for each member of a Joint Venture]].

Date: [insert date (as day, month and year) of Bid Submission]

ICB No.: [insert number of bidding process]

Alternative No.: [insert identification No if this is a Bid for an alternative]

Page ________ of_ ______ pages

1. Bidder’s Name: [insert Bidder’s legal name]

2. Bidder’s JV Member’s name: [insert JV’s Member legal name]

3. Bidder’s JV Member’s country of registration: [insert JV’s Member country of registration]

4. Bidder’s JV Member’s year of registration: [insert JV’s Member year of registration]

5. Bidder’s JV Member’s legal address in country of registration: [insert JV’s Member legal address in country of registration]

6. Bidder’s JV Member’s authorized representative information Name: [insert name of JV’s Member authorized representative] Address: [insert address of JV’s Member authorized representative] Telephone/Fax numbers: [insert telephone/fax numbers of JV’s Member authorized representative] Email Address: [insert email address of JV’s Member authorized representative]

7. Attached are copies of original documents of[check the box(es) of the attached original documents] Articles of Incorporation (or equivalent documents of constitution or association), and/orregistration documents of thelegal entity named above, in accordance with ITB 4.3. In case of a Government-owned enterprise or institution, documents establishing legal and financial autonomy, operation in accordance with commercial law, and absence of dependent status, in accordance with ITB 4.5. 2. Included are the organizational chart, a list of Board of Directors, and the beneficial ownership.

# Price Schedule Forms

[The Bidder shall fill in these Price Schedule Forms in accordance with the instructions indicated.  The list of line items in column 1 of the Price Schedules shall coincide with the List of Goods and Related Services specified by the Purchaser in the Schedule of Requirements.]

| Price Schedule: Goods Manufactured Outside the Purchaser’s Country, to be Imported |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | (Group C bids, goods to be imported) Currencies in accordance with ITB 15 |  |  |  | Date:_________________________ ICB No: _____________________ Alternative No: ________________ Page N ______ of ______ |  |  |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  | 9 |
| Line Item N | Description of Goods | Country of Origin | Delivery Date as defined by Incoterms | Quantity and physical unit | Unit price cip[insert place of destination] in accordance with ITB 14.8(b)(i) | CIP Price per line item (Col. 5x6) | Price per line item for inland transportation and other services required in the Purchaser’s country to convey the Goods to their final destination specified in BDS |  |  | Total Price per Line item (Col. 7+8) |
| [insert number of the item] | [insert name of good] | [insert country of origin of the Good] | [insert quoted Delivery Date] | [insert number of units to be supplied and name of the physical unit] | [insert unit price CIP per unit] | [insert total CIP price per line item] | [insert the corresponding price per line item] |  |  | [insert total price of the line item, only foreign currency] |
|  |  |  |  |  |  |  |  | Total Price |  |  |
| Name of Bidder [insert complete name of Bidder] Signature of Bidder [signature of person signing the Bid] Date [Insert Date] |  |  |  |  |  |  |  |  |  |  |

| Price Schedule: Goods Manufactured Outside the Purchaser’s Country, already imported* |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | (Group C bids, Goods already imported) Currencies in accordance with ITB 15 |  |  |  |  |  | Date:_________________________ ICB No: _____________________ Alternative No: ________________ Page N ______ of ______ |  |  |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 |
| Line Item N | Description of Goods | Country of Origin | Delivery Date as defined by Incoterms | Quantity and physical unit | Unit price including Custom Duties and Import Taxes paid, in accordance with ITB 14.8(c)(i) | Custom Duties and Import Taxes paid per unit in accordance with ITB 14.8(c)(ii) , [to be supported by documents] | Unit Price net of custom duties and import taxes, in accordance with ITB 148 (c) (iii) (Col. 6 minus Col.7) | Price per line item net of Custom Duties and Import Taxes paid, in accordance with ITB 14.8(c)(i) (Col. 58) | Price per line item for inland transportation and other services required in the Purchaser’s country to convey the goods to their final destination, as specified in BDS in accordance with ITB 14.8 (c)(v) | Sales and other taxes paid or payable per item if Contract is awarded (in accordance with ITB 14.8(c)(iv) | Total Price per line item (Col. 9+10) |
| [insert number of the item] | [insert name of Goods] | [insert country of origin of the Good] | [insert quoted Delivery Date] | [insert number of units to be supplied and name of the physical unit] | [insert unit price per unit] | [insert custom duties and taxes paid per unit] | [insert unit price net of custom duties and import taxes] | [ insert price per line item net of custom duties and import taxes] | [insert price per line item for inland transportation and other services required in the Purchaser’s country] | [insert sales and other taxes payable per item if Contract is awarded] | [insert total price per line item] |
|  |  |  |  |  |  |  |  |  |  | Total Bid Price |  |
| Name of Bidder [insert complete name of Bidder] Signature of Bidder [signature of person signing the Bid] Date [insert date] |  |  |  |  |  |  |  |  |  |  |  |

* [For previously imported Goods, the quoted price shall be distinguishable from the original import value of these Goods declared to customs and shall include any rebate or mark-up of the local agent or representative and all local costs except import duties and taxes, which have been and/or have to be paid by the Purchaser. For clarity the bidders are asked to quote the price including import duties, and additionally to provide the import duties and the price net of import duties which is the difference of those values.]

Price Schedule: Goods Manufactured in the Purchaser’s Country

| Purchaser’s Country ______________________ |  |  |  | (Group A and B bids) Currencies in accordance with ITB 15 |  |  |  | Date:_________________________ ICB No: _____________________ Alternative No: ________________ Page N ______ of ______ |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 |
| Line Item N | Description of Goods | Delivery Date as defined by Incoterms | Quantity and physical unit | Unit price EXW | Total EXWprice per line item (Col. 45) | Price per line item for inland transportation and other services required in the Purchaser’s Country to convey the Goods to their final destination | Cost of local labor, raw materials and components from with origin in the Purchaser’s Country % of Col. 5 | Sales and other taxes payable per line item if Contract is awarded (in accordance with ITB 14.8(a)(ii) | Total Price per line item (Col. 6+7) |
| [insert number of the item] | [insert name of Good] | [insert quoted Delivery Date] | [insert number of units to be supplied and name of the physical unit] | [insert EXW unit price] | [insert total EXW price per line item] | [insert the corresponding price per line item] | [Insert cost of local labor, raw material and components from within the Purchase’s country as a % of the EXW price per line item] | [insert sales and other taxes payable per line item if Contract is awarded] | [insert total price per item] |
|  |  |  |  |  |  |  |  | Total Price |  |
| Name of Bidder [insert complete name of Bidder] Signature of Bidder [signature of person signing the Bid] Date [insert date] |  |  |  |  |  |  |  |  |  |

| Price and Completion Schedule - Related Services |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Currencies in accordance with ITB 15 |  |  |  | Date:_________________________ ICB No: _____________________ Alternative No: ________________ Page N ______ of ______ |  |
| 1 | 2 |  | 3 | 4 | 5 | 6 | 7 |
| Service N | Description of Services (excludes inland transportation and other services required in the Purchaser’s country to convey the goods to their final destination) |  | Country of Origin | Delivery Date at place of Final destination | Quantity and physical unit | Unit price | Total Price per Service (Col. 5*6 or estimate) |
| [insert number of the Service ] | [insert name of Services] |  | [insert country of origin of the Services] | [insert delivery date at place of final destination per Service] | [insert number of units to be supplied and name of the physical unit] | [insert unit price per item] | [insert total price per item] |
|  |  |  |  |  | Total Bid Price |  |  |
| Name of Bidder [insert complete name of Bidder] Signature of Bidder [signature of person signing the Bid] Date [insert date] |  |  |  |  |  |  |  |

Form of Bid Security

(Bank Guarantee)

[The bank shall fill in this Bank Guarantee Form in accordance with the instructions indicated.]

[Guarantor letterhead or SWIFT identifier code]

Beneficiary:  [Purchaser to insert its name and address]

IFB No.:  [Purchaser to insert reference number for the Invitation for Bids]

Alternative No.: [Insert identification No if this is a Bid for an alternative]

Date:[Insert date of issue]

BID GUARANTEE No.:[Insert guarantee reference number]

Guarantor:  [Insert name and address of place of issue, unless indicated in the letterhead]

We have been informed that ______ [insert name of the Bidder, which in the case of a joint venture shall be the name of the joint venture (whether legally constituted or prospective) or the names of all members thereof] (hereinafter called "the Applicant") has submitted or will submit to the Beneficiary its bid (hereinafter called "the Bid") for the execution of ________________ under Invitation for Bids No. ___________  (“the IFB”).

Furthermore, we understand that, according to the Beneficiary’s conditions, bids must be supported by a bid guarantee.

At the request of the Applicant, we, as Guarantor, hereby irrevocably undertake to pay the Beneficiary any sum or sums not exceeding in total an amount of ___________  (____________) upon receipt by us of the Beneficiary’s complying demand, supported by the Beneficiary’s statement, whether in the demand itself or a separate signed document accompanying or identifying the demand, stating that either the Applicant:

(a) 	has withdrawn its Bid during the period of bid validity set forth in the Applicant’s  Letter of Bid (“the Bid Validity Period”), or any extension thereto provided by the Applicant; or

(b) 	having been notified of the acceptance of its Bid by the Beneficiary during the Bid Validity Period or any extension thereto provided by the Applicant, (i) has failed to execute the contract agreement, or (ii) has failed to furnish the performance security, in accordance with the Instructions to Bidders (“ITB”) of the Beneficiary’s bidding document.

This guarantee will expire: (a) if the Applicant is the successful bidder, upon our receipt of copies of the contract agreement signed by the Applicant and the performance security issued to the Beneficiary in relation to such contract agreement; or (b) if the Applicant is not the successful bidder, upon the earlier of (i) our receipt of a copy of the Beneficiary’s notification to the Applicant of the results of the bidding process; or (ii)twenty-eight days after the end of the Bid Validity Period.

Consequently, any demand for payment under this guarantee must be received by us at the office indicated above on or before that date.

This guarantee is subject to the Uniform Rules for Demand Guarantees (URDG) 2010 Revision, ICC Publication No. 758.

_____________________________

[Signature(s)]

Note:  All italicized text is for use in preparing this form and shall be deleted from the final product.

Form of Bid Security (Bid Bond)

[The Surety shall fill in this Bid Bond Form in accordance with the instructions indicated.]

BOND NO. ______________________

BY THIS BOND [name of Bidder] as Principal (hereinafter called “the Principal”), and [name, legal title, and address of surety],authorized to transact business in [name of country of Purchaser], as Surety (hereinafter called “the Surety”), are held and firmly bound unto [name of Purchaser] as Obligee (hereinafter called “the Purchaser”) in the sum of [amount of Bond][amount in words], for the payment of which sum, well and truly to be made, we, the said Principal and Surety, bind ourselves, our successors and assigns, jointly and severally, firmly by these presents.

WHEREAS the Principal has submitted or will submit a written Bid to the Purchaser dated the ___ day of ______, 20__, for the supply of [name of Contract] (hereinafter called the “Bid”).

NOW, THEREFORE, THE CONDITION OF THIS OBLIGATION is such that if the Principal:

- has withdrawn its Bid during the period of bid validity set forth in the Principal’s Letter of Bid (“the Bid Validity Period”), or any extension thereto provided by the Principal; or

- having been notified of the acceptance of its Bid by the Purchaser during the  Bid Validity Period or any extension thereto provided by the Principal; (i) failed  to execute the contract agreement; or (ii) has failed to furnish the Performance Security,  in accordance with the Instructions to Bidders (“ITB”) of the Purchaser’s bidding document.

then the Surety undertakes to immediately pay to the Purchaser up to the above amount upon receipt of the Purchaser’s first written demand, without the Purchaser having to substantiate its demand, provided that in its demand the Purchaser shall state that the demand arises from the occurrence of any of the above events, specifying which event(s) has occurred.

The Surety hereby agrees that its obligation will remain in full force and effect up to and including the date 28 days after the date of expiration of the Bid Validity Period set forth in the Principal’s Letter of Bid or any extension thereto provided by the Principal.

IN TESTIMONY WHEREOF, the Principal and the Surety have caused these presents to be executed in their respective names this ____ day of ____________ 20__.

Principal: _______________________	Surety: _____________________________
	Corporate Seal (where appropriate)

_______________________________	____________________________________
(Signature)	(Signature)
(Printed name and title)	(Printed name and title)

Manufacturer’s Authorization

[The Bidder shall require the Manufacturer to fill in this Form in accordance with the instructions indicated. Thisletter of authorization should be on the letterhead of the Manufacturer and should be signed by a person with the proper authority to sign documents that are binding on the Manufacturer.  The Bidder shall include it in its bid, if so indicated in the BDS.]

Date: [insert date (as day, month and year) of Bid Submission]

ICB No.: [insert number of bidding process]

Alternative No.: [insert identification No if this is a Bid for an alternative]

To:  [insert complete name of Purchaser]

WHEREAS

We [insert complete name of Manufacturer], who are official manufacturers of[insert type of goods manufactured], having factories at [insert full address of Manufacturer’s factories], do hereby authorize [insert complete name of Bidder] to submit a bid the purpose of which is to provide the following Goods, manufactured by us [insert name and or brief description of the Goods], and to subsequently negotiate and sign the Contract.

We hereby extend our full guarantee and warranty in accordance with Clause 28 of the General Conditions of Contract, with respect to the Goods offered by the above firm.

Signed: [insert signature(s) of authorized representative(s) of the Manufacturer]

Name: [insert complete name(s) of authorized representative(s) of the Manufacturer]

Title: [insert title]

Dated on ____________ day of __________________, _______ [insert date of signing]

Section V.  Eligible Countries

Eligibility for the Provision of Goods, Works and Non Consulting Services in 
Bank-Financed Procurement

In reference to ITB 4.7 and 5.1, for the information of the Bidders, at the present time firms, goods and services from the following countries are excluded from this bidding process:

Under ITB 4.7(a) and 5.1:	 [insert a list of the countries following approval by the Bank to apply the restriction or state “none”].

Under ITB 4.7(b) and 5.1:	 [insert a list of the countries following approval by the Bank to apply the restriction or state “none”]

Section VI. Bank Policy - Corrupt and Fraudulent Practices

Guidelines for Procurement of Goods, Works, and Non-Consulting Services under IBRD Loans and IDA Credits & Grants by World Bank Borrowers, dated January 2011.

“Fraud and Corruption:

1.16	It is the Bank’s policy to require that Borrowers (including beneficiaries of Bank loans), bidders, suppliers, contractors and their agents (whether declared or not), sub-contractors, sub-consultants, service providers or suppliers, and any personnel thereof, observe the highest standard of ethics during the procurement and execution of Bank-financed contracts. In pursuance of this policy, the Bank:

(a)	defines, for the purposes of this provision, the terms set forth below as follows:

(i)	“corrupt practice” is the offering, giving, receiving, or soliciting, directly or indirectly, of anything of value to influence improperly the actions of another party;;

(ii) 	“fraudulent practice” is any act or omission, including a misrepresentation, that knowingly or recklessly misleads, or attempts to mislead, a party to obtain a financial or other benefit or to avoid an obligation;

(iii)	“collusive practice” is an arrangement between two or more parties designed to achieve an improper purpose, including to influence improperly the actions of another party;

(iv)	“coercive practice” is impairing or harming, or threatening to impair or harm, directly or indirectly, any party or the property of the party to influence improperly the actions of a party;

(v)	"obstructive practice" is:

(aa)	deliberately destroying, falsifying, altering, or concealing of evidence material to the investigation or making false statements to investigators in order to materially impede a Bank investigation into allegations of a corrupt, fraudulent, coercive or collusive practice; and/or threatening, harassing or intimidating any party to prevent it from disclosing its knowledge of matters relevant to the investigation or from pursuing the investigation, or

(bb)	acts intended to materially impede the exercise of the Bank’s inspection and audit rights provided for under paragraph 1.16(e) below.

(b)	will reject a proposal for award if it determines that the bidder recommended for award, or any of its personnel, or its agents, or its sub-consultants, sub-contractors, service providers, suppliers and/or their employees, has, directly or indirectly, engaged in corrupt, fraudulent, collusive, coercive, or obstructive practices in competing for the contract in question;

(c)	will declare misprocurement and cancel the portion of the loan allocated to a contract if it determines at any time that representatives of the Borrower or of a recipient of any part of the proceeds of the loan engaged in corrupt, fraudulent, collusive, coercive, or obstructive practices during the procurement or the implementation of the contract in question, without the Borrower having taken timely and appropriate action satisfactory to the Bank to address such practices when they occur, including by failing to inform the Bank in a timely manner at the time they knew of the practices;

(d)	will sanction a firm or individual, at any time, in accordance with the prevailing Bank’s sanctions procedures, including by publicly declaring such firm or individual ineligible, either indefinitely or for a stated period of time: (i) to be awarded a Bank-financed contract; and (ii) to be a nominated;

(e)	will require that a clause be included in bidding documents and in contracts financed by a Bank loan, requiring bidders, suppliers and contractors, and their sub-contractors, agents, personnel, consultants, service providers, or suppliers, to permit the Bank to inspect all accounts, records, and other documents relating to the submission of bids and contract performance, and to have them audited by auditors appointed by the Bank.”

# PART 2 – Supply Requirements

Section VII. Schedule of Requirements

Contents

1.  List of Goods and Delivery Schedule	65

2.	List of  Related Services and Completion Schedule 	89

3.	Technical Specifications	91

4.	Drawings	177

5. 	Inspections and Tests	178

| 1. List of Goods and Delivery Schedule Lot 1- Diagnostic and Imaging equipment |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Line Item N | Description of Goods | Quantity | Physical unit | Final (Project Site) Destination as specified in BDS | Delivery (as per Incoterms) Date |  |  |
|  |  |  |  |  | Earliest Delivery Date | Latest Delivery Date | Bidder’s offered Delivery date [to be provided by the bidder] |
|  | Digital Universal X-Ray Unit | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | Not acceptable | 120 days after Contract signature | [insert the number of days following the date of Contract signature] |
|  | CT scanner 32 slice | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Mobile X-Ray | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | UPS for CT scanner, digital universal X-ray and digital angiography system | 2 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Surgical X-Ray unit (C-Arm) | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Tabletop dry imager for CT scanner, digital universal X-ray | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | X-ray Film Viewer | 10 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Ultrasound (3 probes) with printer | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Portable ultrasound | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Floor-mounted digital angiography system for interventional cardiology | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |

| List of Goods and Delivery Schedule Lot 2- Medical Laboratory equipment |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Line Item N | Description of Goods | Quantity | Physical unit | Final (Project Site) Destination as specified in BDS | Delivery (as per Incoterms) Date |  |  |
|  |  |  |  |  | Earliest Delivery Date | Latest Delivery Date | Bidder’s offered Delivery date [to be provided by the bidder] |
|  | Hematological analyzer | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature | [insert the number of days following the date of Contract signature] |
|  | Semi-Automatic Chemistry Analyzer | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Coagulometer | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Critical Care Analyser | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Laboratory scale | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Water Bath | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Water Still | 2 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Table Top Centrifuge | 4 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Microscope Binocular | 4 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Hot air sterilizer | 6 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Incubator | 2 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Autoclave Table-top | 2 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Autoclave Steam Sterilizer | 2 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Laboratory instrument set | 2 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | ELISA (reader, washer, shaker) | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Urinalysis System | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |

| List of Goods and Delivery Schedule Lot 3- Surgery and Hospital equipment |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Line Item N | Description of Goods | Quantity | Physical unit | Final (Project Site) Destination as specified in BDS | Delivery (as per Incoterms) Date |  |  |
|  |  |  |  |  | Earliest Delivery Date | Latest Delivery Date | Bidder’s offered Delivery date [to be provided by the bidder] |
|  | Electrocardiograph | 3 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature | [insert the number of days following the date of Contract signature] |
|  | Stress-Test system | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | OT Light (LED type) | 4 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | OT Light (LED type) Mobile OT Light (LED type) | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Operating table multifunctional with orthopedic attachment | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Operating Table Universal | 3 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Anesthesia trolley (anesthesia machine with ventilator and monitor) | 4 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Electrosurgical Unit | 4 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Syringe pump | 7 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Operation instrument set | 4 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Ventilator (Intensive ventilation equipment) | 3 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Defibrillator with monitor | 4 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Volumetric Infusion Pump | 15 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Patient Monitor | 29 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Syringe pump | 15 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Oxygen Concentrator | 4 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Holter Recorder with soft and computer | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Slit lamp | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Ophthalmoscope and otoscope set | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | EMG | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | AMBU emergency case | 2 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | UV-Air Flow Cleaner–Recirculators Model Recirculator | 23 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Nebulizer | 2 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Examination Lamp | 12 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Scale for Adults | 3 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Laparoscopy set | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Head light | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Neurosurgical Instrument set | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Set of instruments for installing intramedullary nails for femoral bone with case (universal) | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Surgical ENT set | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Audiometer | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | CPAP machine | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Newborn suction pump | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Endovenous laser | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Haemodialysis Machine and accessories | 15 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |
|  | Water treatment plant for hemodialysis equipment | 1 | Unit | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build | N/A | 120 days after Contract signature |  |

| List of Goods and Delivery Schedule Lot 4- Medical furniture |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Line Item N | Description of Goods |  | Quantity | Physical unit |  | Final (Project Site) Destination as specified in BDS |  | Delivery (as per Incoterms) Date |  |  |  |  |
|  |  |  |  |  |  |  |  | Earliest Delivery Date | Latest Delivery Date |  | Bidder’s offered Delivery date [to be provided by the bidder] |  |
|  | Reception desk |  | 1 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  | [insert the number of days following the date of Contract signature] |  |
|  | Reception desk for outpatient department |  | 1 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Reception desk desk for emergency |  | 1 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Plastic Chairs |  | 520 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Writing Desk |  | 80 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Cabinet |  | 45 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Chair |  | 90 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Wardrobe |  | 78 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Bed With Backrest |  | 146 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Rehabilitation Bed (electric bed) |  | 35 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Bedside cabinet |  | 193 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Movable table |  | 193 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Bedside Locker |  | 46 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Baby bed |  | 6 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Baby crib |  | 5 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Hanger |  | 75 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Soled linen trolley |  | 6 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Bowl`s stand |  | 6 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Table for laboratory equipment |  | 7 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Medical examination couch |  | 40 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Cabinet for medicine |  | 65 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Color LED TV - 43"with holder |  | 3 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Color LED TV - 32"with holder |  | 78 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Refrigerator |  | 83 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Cabinet with Shelves |  | 75 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Table |  | 100 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Nurse station#1 |  | 8 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Nurse station#2 |  | 6 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Table for medical instruments (B) |  | 15 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Table for medical instruments (A) |  | 35 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Waste bin |  | 190 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Folding screen |  | 23 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Mounting step |  | 12 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Infusion fluid holder |  | 70 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Revolving chair with gas pump |  | 12 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Changing room locker, single door |  | 70 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Kitchen refrigerator |  | 12 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Kitchen sink with electric hod#1 |  | 1 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Kitchen sink with electric hod#2 |  | 1 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Kitchen sink with electric hod#3 |  | 4 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Patient trolley |  | 6 |  |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
|  | Washing machine |  | 5 | Unit |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | N/A | 120 days after Contract signature |  |  |  |
| 2. List of Related Services and Completion Schedule |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Service |  | Description of Service |  |  | Quantity |  | Physical Unit |  |  | Place where Services shall be performed |  | Final Completion Date(s) of Services |  |
| Lot 1- Diagnostic and Imaging equipment |  | Installation of all equipment |  |  | 20 |  | Unit |  |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | within 25 days after delivery to the final destination goods |  |
|  |  | Training |  |  | During 5 days |  | 13 persons (doctors-5; X-ray technicians-7; nurse-1) |  |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | within 15 days after installation |  |
| Lot 2- Medical Laboratory equipment |  | Installation of all equipment |  |  | 32 |  | Unit |  |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | within 25 days after delivery to the final destination goods |  |
|  |  | Training |  |  | During 5 days |  | 21 persons (doctors-5; laboratory technicians-16) |  |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | within 15 days after installation |  |
| Lot 3- Surgery and Hospital equipment |  | Installation of all equipment |  |  | 173 |  | Unit |  |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | within 25 days after delivery to the final destination goods |  |
|  |  | Training |  |  | During 5 days |  | 28 persons (doctors-12; anesthesiology technicians-16) |  |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | within 15 days after installation |  |
| Lot 4- Medical furniture |  | Installation of all furniture |  |  | 2448 |  | Unit |  |  | “Vanadzor MC” CJSC, Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |  | within 25 days after delivery to the final destination goods |  |

3.	Technical Specifications

Lot-1-Diagnostic and Imaging equipment

| N | NAME | TECHNICAL SPECIFICATION | Q-ty |
| --- | --- | --- | --- |
| 1 | Թվային ռենտգեն համալիր ունիվերսալ// Digital Universal X-Ray Unit | Digital radiography and fluoroscopy system | 1 |
|  |  | System specifications |  |
|  |  | Universal Unit Fully digital Remote controlled overtable system for digital fluoroscopy and digital radiography as a 2-in-1 solution. |  |
|  |  | Multifunctional, workflow-optimized system remote control console for controlling system , generator functions. |  |
|  |  | Microprocessor-controlled user interface for generator, imaging chain and image functions, selection of system functions and display of system statuses. Intuitive, ergonomic operation through direct selection of workplaces and for fully digital image chain with network connectivity in DICOM standard: Send, Print and Storage Commitment. DICOM DVD / CD burner. |  |
|  |  | Overhead Tube on the freely moving Tube Arm mounted on the patient Table, without any floor or ceiling mounted column. |  |
|  |  | The System to be suitable for Standard Skeletal and Radiographic Examinations, Including Oblique Beam Proje ctions. |  |
|  |  | Table. |  |
|  |  | Tilt: Motor Driven: from +89º to -89º, deviation not more than 10% automatic stop in horizontal position |  |
|  |  | Height: from 48 to 120 cm (variable), deviation not more than 20cm |  |
|  |  | Tabletop Outside Dimensions: |  |
|  |  | Length: 225 cm, deviation not more than 10cm |  |
|  |  | Width: not less than 73cm |  |
|  |  | Radiolucent (X/Y): not less than 190 cm x 50 cm |  |
|  |  | Traversal tabletop movement: from +15cm to -15cm, deviation not more than 10% |  |
|  |  | Patient scan range: 160cm, deviation not more than 20cm |  |
|  |  | Patient weight: at least 200 kg |  |
|  |  | Tube Assembly: |  |
|  |  | Two-focus X-ray tube |  |
|  |  | Max. exposure voltage: 150 kV, deviation not more than 10% |  |
|  |  | Anode heat storage capacity: 600 000 HU, deviation not more than 10% |  |
|  |  | Anodes drive: from 3000 to 10000 U/min, deviation not more than 10% |  |
|  |  | Focal spot nominal value: |  |
|  |  | Small focus: 0.4mm, deviation not more than 0.2mm |  |
|  |  | Large focus: 1.0mm, deviation not more than 0.2mm |  |
|  |  | Nominal output of anode: |  |
|  |  | Small: not less than 23 kW |  |
|  |  | Large: not less than 42 kW |  |
|  |  | Complete filtration: 2.5 mm Al , deviation not more than 10% |  |
|  |  | Oblique Projections: CRA -35º, deviation not more than 10o / CAU+30º, deviation not more than 10o |  |
|  |  | Flat panel detector |  |
|  |  | Will be present the digital high-resolution dynamic flat detector with integrated removable grid |  |
|  |  | Detector size: 43cmx43cm, deviation not more than 3% |  |
|  |  | Pixel size: not more than 200 μm |  |
|  |  | Image matrix size: not less than 1024x1024 pixel |  |
|  |  | Detail resolution: not less than 2.5 PL/mm |  |
|  |  | Usable input formats: |  |
|  |  | Overview: not less than 42cmx42cm |  |
|  |  | Zoom 1 not less than 30cm x30 cm |  |
|  |  | Zoom 2 not less than 20cm x 20cm |  |
|  |  | Zoom 3 not less than 15cm x 15cm |  |
|  |  | Operating Modes |  |
|  |  | Digital ragiography |  |
|  |  | Single image and serial mode not less than 1024 x 1024 12-bit matrix |  |
|  |  | Serial mode with variable frame rate max. 15 f/s, deviation not more than 10% |  |
|  |  | Last Image Hold (LIH) |  |
|  |  | Measuring chamber for acquisition of the dose-area product and display of the value for measured dose-area product is and the calculated patient entry dose |  |
|  |  | Radiation free placement of collimator blades in the LIH |  |
|  |  | Digital Density Optimization |  |
|  |  | Image Processing |  |
|  |  | Real-time edge enhancement |  |
|  |  | positive/negative image display |  |
|  |  | electronic display (shutter) |  |
|  |  | vertical and horizontal image inversion |  |
|  |  | magnifying glass functions |  |
|  |  | zoom functions |  |
|  |  | post-processing with high detail contrast and reduced noise |  |
|  |  | quantification: angle/length measurement |  |
|  |  | Image display |  |
|  |  | Image review and display in 100 Hz progressive display (1024 x 1024 matrix) through high-resolution, flicker- free flat-screen displays |  |
|  |  | Screen layout with 4, 9, 16, 64 images of an examination |  |
|  |  | Image storage capacity |  |
|  |  | No. of images for permanent storage in 1k/14-bit matrix, 70,000, deviation not more than 10% |  |
|  |  | No. of images for permanent storage in 2840 x 2880 matrix not less than 2,000 |  |
|  |  | Compression unit |  |
|  |  | Remote-controlled compression device integrated in the tube support |  |
|  |  | LCD display |  |
|  |  | Mobile display trolley and remote-control of table with 19" medical grade monitor for examination room |  |
|  |  | Medical grade monitor for live image display in the control room: 19" medical grade monitor. Monochrome flat-screen 21" system display in control room (system monitor) |  |
|  |  | Connectivity |  |
|  |  | DICOM Send |  |
|  |  | DICOM Print |  |
|  |  | DICOM Image Storage |  |
|  |  | DVD for Fluoro Recording |  |
|  |  | Power Connection |  |
|  |  | Nominal Voltage: 3/N/PE, 380V |  |
|  |  | Nominal Frequency: 50Hz. |  |
|  |  | Completeness: |  |
|  |  | Lead glass, radiation protection window 90x100cm., 2.0mm. Pb, deviation not more than 10% |  |
|  |  | Accessories: |  |
|  |  | The set includes all the necessary additional devices and accessories to complete functioning of the equipment (Example: Dosimeter, cable, powerbox, installation material etc.) |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |
|  |  | ISO13485 or equivalent |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |
| 2 | Համակարգչային շերտագիր տոմոգրաֆ (32 slice) // CT scanner 32 slice | Multislice CT scanner adapted specifically for clinical routine. | 1 |
|  |  | The simplified user interface guides the operator, and wherever possible automated workflow. |  |
|  |  | Gantry |  |
|  |  | Aperture [cm]: not less than 65 cm |  |
|  |  | Slip ring for power supply [low voltage] |  |
|  |  | Tilt range [degree]: not less than from 30 to -30 , increment not more than 0.5 or digital tilt |  |
|  |  | Positioning lights - panels (Transaxial, sagittal, coronal) |  |
|  |  | Positioning lights type (Laser) |  |
|  |  | Positioning lights accuracy(at centre of gantry) [mm]: from 1 to -1 |  |
|  |  | Gantry weight [kg]: not more than 1400 |  |
|  |  | X-ray generator |  |
|  |  | Location: Rotation assembly or Gantry |  |
|  |  | Method of generator cooling: Air or Water |  |
|  |  | Power rating [kW]: not more than 100 |  |
|  |  | kV settings ranges: from 80 to 130 |  |
|  |  | mA range: from 10 to 200, increment not more than 5mA step |  |
|  |  | X-ray tube |  |
|  |  | Anode heat capacity [MHU]: not less than 2 |  |
|  |  | Maximum anode cooling rate [kHU/min]: not less than 500 |  |
|  |  | Method of tube cooling: Oil or Air |  |
|  |  | Expected tube life: not less than 6000 exams or 3 years |  |
|  |  | Detection system |  |
|  |  | Detector type (Solid state array) |  |
|  |  | Maximum number of slices: not less than 32 |  |
|  |  | Hardware features to improve xy-plane resolution |  |
|  |  | Hardware features to improve z-axis resolution |  |
|  |  | High contrast spatial resolution [Ipcm]: not less than 15.4 |  |
|  |  | Couch |  |
|  |  | Horizontal movement range [cm]: not less than 152cm |  |
|  |  | Couch control on ganrty (Front, Left, Right) |  |
|  |  | Couch control in room |  |
|  |  | Couch control on console |  |
|  |  | Scanned projection radiography (SPR) |  |
|  |  | SPR dimensions [mm x mm]: adjustable |  |
|  |  | Angular positions of X-ray tube availability for SPR [degrees]: from 0 to 359, increment not more than 1 degree step or lateral |  |
|  |  | Maximum number of SPR viewsavailability for scan planing: 2 |  |
|  |  | Accuracy of slice prescription from SPR [mm]: not more than ±0.5 |  |
|  |  | Parameters |  |
|  |  | Minimum reconstruction field of view: from 5 to 9.6 cm |  |
|  |  | Maximum reconstruction field of view at full image quality: from 43to 50 cm |  |
|  |  | Reconstruction matrices: not less than 512 x 512 |  |
|  |  | Sequential image slice option |  |
|  |  | Beam collimation slice widths (mm): 0.625, 1.25, 2.5, 5, 10 or equivalent |  |
|  |  | Helical image slice option |  |
|  |  | Beam collimation slice widths (mm): 0.625, 1.25, 2.5, 5, 10 or equivalent |  |
|  |  | Minimum reconstruction interval of images, helical scanning (mm): 0.1 |  |
|  |  | Functionality |  |
|  |  | Number of prospective reconstructions that can be set up in protocol |  |
|  |  | Simultaneous scanning and routine analysis |  |
|  |  | Simultaneous scanning and archiving and/or hard copy |  |
|  |  | Data processing |  |
|  |  | Reconstruction algorithm - sequential scanning: Filtered back projection and iterative reconstruction |  |
|  |  | Reconstruction approach - sequential scanning: minimum 2D |  |
|  |  | Cone beam reconstruction - sequential scanning |  |
|  |  | Cone beam reconstruction - helical scanning |  |
|  |  | Low signal correction |  |
|  |  | Motion correction |  |
|  |  | Beam hardening algorithms |  |
|  |  | Noise reduction techniques - raw data based |  |
|  |  | Other artefact reduction methods |  |
|  |  | System start up and calibration |  |
|  |  | Total time from standby mode to scanning (mins): not more than 30 |  |
|  |  | Total time from standby mode to scanning in an emergency (mins): not more than 1 |  |
|  |  | Power on to ready to warm up (mins): not more than 10 |  |
|  |  | Standby to ready to warm up: not more than 1 |  |
|  |  | Tube warm up time (mins): not more than 5 |  |
|  |  | Performance data |  |
|  |  | In plane spatial resolution for sharpest clinical algorithm (lp/cm) (not less than or equivalent): |  |
|  |  | MTF0: 15.4 |  |
|  |  | MTF10: 13.0 |  |
|  |  | MTF50: 8.5 |  |
|  |  | Longitudinal (z -axis) spatial resolution for sharpest clinical algorithm (lp/cm) (not less than or equivalent): |  |
|  |  | MTF0: 15.4 |  |
|  |  | MTF10: 13.0 |  |
|  |  | MTF50: 8.5 |  |
|  |  | Contrast resolution: smallest rod size discernable (in plane) (mm): from 3 to 5 |  |
|  |  | Workstation with all necessary equipment and software |  |
|  |  | Input device: monitor, keyboard, mouse |  |
|  |  | Number of monitor: not less than 2 |  |
|  |  | Display: LCD type not less than 19" |  |
|  |  | Main computers |  |
|  |  | Reconscructor |  |
|  |  | Max.no. of slices displayed at once: 32 |  |
|  |  | Storage (at scanner) |  |
|  |  | Image archive |  |
|  |  | A4 Laser printer |  |
|  |  | Laser printer for CT scanner |  |
|  |  | Software |  |
|  |  | CT angiography basic package |  |
|  |  | Brain perfusion software |  |
|  |  | Body perfusion software |  |
|  |  | Computer assisted reading (Lung) |  |
|  |  | Computer assisted reading (Colon) |  |
|  |  | Bone mineral densitomentry |  |
|  |  | Dose modulation technique |  |
|  |  | Accessories: |  |
|  |  | The set includes all the necessary additional devices, accessories to complete functioning of the equipment and safety accessories required for the protection of personnel (example: workstation, contrast injector with 2 syringes, distance remote control and warm function, headrests, armrests, holders cover glass, cable, installation materials,insulation materials and ect) |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |
|  |  | ISO13485 or equivalent |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |
| 3 | Շարժական ռենտգեն համակարգ // Mobile X-Ray | For transported from department to department, moving in elevators, over thresholds and around corners. For use in the intensive care, premature children, accident and X-ray wards. | 1 |
|  |  | Mobile radiographic unit: not less than 15kW |  |
|  |  | Power requirements: automatic line voltage adaptation 220V/ 50Hz |  |
|  |  | Shoul have backup battery |  |
|  |  | High-frequency generator with single-tank tube unit, multipulse voltage form |  |
|  |  | kV range: from 40 kV to 125 kV, deviation not more than 10% |  |
|  |  | mAs range: from 0.2 to 220 mAs, deviation not more than 10% |  |
|  |  | Maximum switching time: not more than 8s |  |
|  |  | Remote control |  |
|  |  | Counterbalanced tube support arm |  |
|  |  | Exact positioning of the unit by swiveling castors |  |
|  |  | Single-tank tube unit with built-in rotating anode tube: 3000 rpm (50 Hz), |  |
|  |  | Anode thermal capacity: not less than 107000 HU |  |
|  |  | Focal spot from: min from 0.6 to 0.8 to max from 1.0 to 1.3mm |  |
|  |  | Anodic angle: 15o |  |
|  |  | Manual collimator with internal light source, multilayer, square field |  |
|  |  | Collimator rotation: min. from 90o to -90o |  |
|  |  | Cassette compartment for holding: not less than 4 cassettes, 35x43cm (14x17") |  |
|  |  | Should have cassette 4 different size |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |
|  |  | ISO13485 or equivalent |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |
|  |  | X-Ray Film processor tabletop |  |
|  |  | Processing machine for X-ray films from: |  |
|  |  | from 10x10cm to 35x43cm |  |
|  |  | Processing chemicals Commercial products suitable for table-top machines |  |
|  |  | Developer, Fixer, Watertank capacity: not less than 2.1L |  |
|  |  | Replenishment bottles: not less than 5L each |  |
|  |  | Heating time not more than 10min. |  |
|  |  | Films All double-sided coated general X-ray films suitable for automatic processing. |  |
|  |  | Film throughput not less than 50 pcs./hour for 35cmx43cm cassette |  |
|  |  | Film throughput speed: not less than 35cm /min |  |
|  |  | Processing time: not more than 100 sec |  |
|  |  | Noise level: not more than 60dB (A) |  |
|  |  | Power supply: 220V/50Hz |  |
|  |  | Max. power consumption: not more than 1200 W |  |
|  |  | Dimensions: not more than 90L x 70W x 45H cm |  |
|  |  | Replenishment bottles are supplied with machine |  |
|  |  | Accessories |  |
|  |  | Light-tight cover |  |
|  |  | Permanent water connection |  |
|  |  | External replenishment system |  |
|  |  | Table for X-Ray Film processor |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |
|  |  | ISO13485 or equivalent |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |
| 4 | Անխափան սնուցման համակարգ // UPS for CT scanner, digital universal X-ray and digital angiography system | UPS (UNINTERRUPTIBLE POWER SUPPLY ) for CT scanner and digital universal X-ray not less than 120kVA | 2 |
|  |  | From Critical Power Supplies must be an on-line double-conversion UPS and designed to give not less than 20 minutes run time, of simultaneous work CT scanner and digital X-ray and digital angiography system without power on input. |  |
|  |  | The UPS has a three phase input and output. |  |
|  |  | Technology: On-line UPS |  |
|  |  | Format: Floor Standing |  |
|  |  | Battery unit must be designed for the special battery and is made by the same manufacturer as the electronic control unit. |  |
|  |  | Gel batteries should be |  |
|  |  | Accessories |  |
|  |  | The set includes all the necessary additional devices, accessories to complete functioning of the equipment (example: air conditioning, cable, installation materials and ect) |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |
|  |  | ISO13485 or equivalent |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |
| 5 | Վիրաբուժական ռենտգեն համակարգ (ЭОП) // Surgical X-Ray unit (C-Arm) | Operating modes | 1 |
|  |  | Fluoroscopy with Last Image Hold |  |
|  |  | Pulsed Fluoroscopy with Last Image Hold |  |
|  |  | Digital Radiography |  |
|  |  | Digital Angiography |  |
|  |  | Digital Substraction Angiography |  |
|  |  | Selection from 5 application-specific fluoroscopy and radiography curves. |  |
|  |  | C-Arm |  |
|  |  | Orbital movement: not less than 120° |  |
|  |  | Angulation: not less than from 190° to -190° |  |
|  |  | Horizontal movement: 20 cm, deviation not more than 10% |  |
|  |  | Swivel range: not less than 10° to -10° |  |
|  |  | Vertical Travel: 40 cm, deviation not more than 10% |  |
|  |  | Source – I.I. distance: 80 cm, deviation not more than 10% |  |
|  |  | Immersion depth 60 cm, deviation not more than 10% |  |
|  |  | Lateral movement: steering wheel |  |
|  |  | X-ray generator |  |
|  |  | Single-tank high-frequency converter control not less than 15kHz |  |
|  |  | kV range: from 40kV to 110kV |  |
|  |  | Nominal power output: not more than 10 kW |  |
|  |  | Radiography |  |
|  |  | Digital Radiography (mA): from 0.2mA to 20mA |  |
|  |  | Cassette exposures max: not more than 150mA |  |
|  |  | Fluoroscopy |  |
|  |  | mA: from 0.2mA to 8.0mA or more |  |
|  |  | X-ray tube unit |  |
|  |  | Should have two focal spots |  |
|  |  | Small: not more than 0,5 mm |  |
|  |  | Large: not more than 1.5 mm |  |
|  |  | Nominal voltage: 110kV, deviation not more than 10% |  |
|  |  | Collimator |  |
|  |  | Iris collimator |  |
|  |  | Cassette holder |  |
|  |  | Format 24x30cm or equivalent size |  |
|  |  | Nominal diameter: 23cm, deviation not more than 10% |  |
|  |  | Zoom: 15cm, deviation not more than 10% |  |
|  |  | Metal-enamel technology, anti-glare output screen, high-transparency input window or antiscatter grid 8:1 |  |
|  |  | TV system |  |
|  |  | DC with CCD sensor |  |
|  |  | A/D conversion: not more than 14 bit |  |
|  |  | Resolution: not less than 1024x1024 |  |
|  |  | Image continuous rotation |  |
|  |  | Monitor |  |
|  |  | LCD displays not less than 19" high-resolution, high-brightness, anti-glare output screen |  |
|  |  | Display: not less than 2 |  |
|  |  | Maximum brightness: not less than 1000 cd/m2 |  |
|  |  | Digital image storage |  |
|  |  | Storage capacity: not less than 15000 images or 500 Gb |  |
|  |  | Image processing |  |
|  |  | Noise reduction; frequency filtration; rotation; positive/negative inversion; |  |
|  |  | left/right and top/bottom reversal; subtraction |  |
|  |  | Power supply |  |
|  |  | Power requirement maximum: not more than 12kVA |  |
|  |  | Main voltage single phase 220V/50 Hz |  |
|  |  | Accessories |  |
|  |  | The set includes all the necessary additional devices, accessories to complete functioning of the equipment (example: sterile cover and clips for C-arm, X-ray tube and I.I and ect) |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |
|  |  | ISO13485 or equivalent |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |
| 6 | Չոր երևակման սարք // Tabletop dry imager for CT scanner, digital universal X-ray | Multi- size direct digital tabletop solid-state technology DICOM-native imager, which makes network connectivity easy. No wet processing, no darkroom, no complicated adjustments or cleaning procedures. | 1 |
|  |  | It designed and can be connected so as to provide both simultaneous print from digital X-ray and CT scanner |  |
|  |  | Sizes: 8 x 10 inch, 10 x 12 inch, 10 x 14 inch, 14 x 17 inch; min. Two on-line sizes |  |
|  |  | Throughput: |  |
|  |  | 8 x 10 inch: not less than 100 sheets/hour |  |
|  |  | 10 x 14 inch: not less than 100 sheets/hour |  |
|  |  | 14 x 17 inch: not less than 80 sheets/hour |  |
|  |  | Printing resolution: Geometrical - 254ppi (100micron pixel) -508 ppi(50 micron pixel); |  |
|  |  | Contrast: 14 bits, deviation not more than 10% |  |
|  |  | Weight (without film): not more than 104 kg |  |
|  |  | Dimensions (W x D x H): not more than 61 x 63 x 89 cm |  |
|  |  | Power consumption: Peak - 1500 Watt, deviation not more than 10%; Printing - 350 Watt, deviation not more than 10%; Standby: 60 Watt, deviation not more than 10% |  |
|  |  | Capacity of supply tray: 100 sheets (size 14x17 inch) - 150 sheet (sizes 8x10, 10x12, 10x14inch) |  |
|  |  | Power supply: 220V/50Hz |  |
|  |  | Power supply: 220V/50Hz |  |
|  |  | Accessories: |  |
|  |  | The set includes all the necessary additional devices and accessories to complete functioning of the equipment |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |
|  |  | ISO13485 or equivalent |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |
| 7 | Նեգատոսկոպ // X-ray Film Viewer | X-ray Film Illuminator (Viewer), for two films 35x43cm. | 10 |
|  |  | Wall-Mount not less than 15-watt Fluorescent tubes with a color Temperature of 6500° Kelvin color, deviation not more than 10% provide a Bright and even Spread of light. |  |
|  |  | Transparent Spring-loaded film Retainers grip lightly and firmly without obscuring top edge details |  |
|  |  | Rigid welded steel construction finished in white enamel with stainless steel trim |  |
|  |  | Screen recessed into the cabinet to help keep the interior dust free and eliminate side light spill |  |
|  |  | Flicker-free Capability |  |
|  |  | The front acrylic Screen not fixed directly to the cabinet and so free to move in order to eliminate screen distortion due to thermal expansion. The film retainers always operate effectively maintaining an even pressure across the full width of the illuminator. |  |
|  |  | Inner Paint: Optical White |  |
|  |  | Direct Starting, Luminous Source: Daylight |  |
|  |  | Front Side: ON/OFF button |  |
|  |  | Bipolar Switch with Pilot lamp |  |
|  |  | Wall Fixations provided for Cord , Local Plug |  |
|  |  | Power Supply: 220V/50Hz |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |
|  |  | ISO13485 or equivalent |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |
| 8 | Ուլտրաձայնային սարք գունավոր դոպլերով 3 տվիչ // Ultrasound (3 probes) with printer | Ultrasound System with 3 probes and printer. | 1 |
|  |  | Applications: Abdominal, Obstetrical, Gynecological, Small Parts, Musculoskeletal, Vascular/ Peripheral Vascular, Urological, Pediatric |  |
|  |  | Scanning methods: |  |
|  |  | Electronic Convex |  |
|  |  | Electronic Linear |  |
|  |  | Electronic Micro Convex |  |
|  |  | Electronic Sector |  |
|  |  | Real Time 4D Volume Sweep |  |
|  |  | Operating modes |  |
|  |  | B-Mode |  |
|  |  | Coded Phase Inversion Harmonic Imaging |  |
|  |  | M-Mode |  |
|  |  | Color M Mode |  |
|  |  | Color Flow Mode (CFM) |  |
|  |  | Power Doppler Imaging (PDI) |  |
|  |  | Directional PDI |  |
|  |  | PW Doppler with High PRF |  |
|  |  | Anatomical M-Mode |  |
|  |  | CW Doppler Mode |  |
|  |  | 3D Volume Modes |  |
|  |  | Display modes |  |
|  |  | Simultaneous Capability: B/PW, B/CFM, B/M, B + CFM/M, Real Time Triplex Mode (B + CFM), Dual B (B/B) |  |
|  |  | Selectable alternating modes: B/M, B/PW, B + CFM/M, B + CFM (PDI)/PW (CW), 3D-Mode, 3D-Mode Color, B/CW, B + CFM/CW, Multi-image split screen (quad screen), Live and/or frozen, B + B/CFM, PDI, PW/M, Independent CINE playback |  |
|  |  | Zoom: Write/Read/Pan |  |
|  |  | Colorized Image |  |
|  |  | Measurement & Analysis mode: |  |
|  |  | On B-mode |  |
|  |  | Distance Area and Circumference by ellipse and trace Angle Volume, Stenotic rate (%STENO DIST, %STENO Area-T) Ratio (Ratio Dist, Ratio Area-T) |  |
|  |  | On M-mode image |  |
|  |  | Velocity Time interval Distance (amplitude) Heart rate Stenotic rate (%STENO Length) |  |
|  |  | Obstetrical Measurements |  |
|  |  | Gynecological Measurements |  |
|  |  | Cardiac Measurements |  |
|  |  | Maximum depth of field: 33 cm, deviation not less than 10% |  |
|  |  | Probe connectors: |  |
|  |  | Active connectors: 3 connectors |  |
|  |  | Viewing Monitor: 19” diagonal color LCD display |  |
|  |  | Touch screen LCD monitor on board operator keyboard: 8.4", deviation not less than 10% |  |
|  |  | Articulating monitor arm: |  |
|  |  | Probes: not less than 3 probes: |  |
|  |  | Electronic Convex Sector: from 2 to 5Mhz or equivalent |  |
|  |  | Electronic Linear: from 4 to 13Mhz or equivalent |  |
|  |  | Electronic Phased Array Sector: from 1.7 to 4Mhz or equivalent |  |
|  |  | Should have optional support of 4D volume probes |  |
|  |  | Power Requirements: 220V / 50 Hz. |  |
|  |  | Accessories: |  |
|  |  | Ultrasound gel - 1bottle, printer - 1pc, paper for printer -2rol |  |
|  |  | On line UPS: not less than 20 minutes run time |  |
|  |  | The set includes all the necessary additional devices and accessories to complete functioning of the equipment |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |
|  |  | ISO13485 or equivalent |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |
| 9 | Շարժական ուլտրաձայնային սարք // Portable ultrasound | Portable ultrasound system of advanced technologies such as Color, Power, and Spectral Doppler, Harmonic Imaging, Speckle Reduction, Spatial Compounding, Trapezoid Imaging, and Panoramic Imaging with 2 probes | 1 |
|  |  | Applications: Abdominal (Adult, Pediatric, Neonatal), Adult Transcranial, Breast, Cardiac (Adult, Pediatric, Neonatal), Emergency, General Imaging, Musculoskeletal, OB/GYN, Small Parts, and Vascular. |  |
|  |  | Display Screen:Color LCD: not less than 15" |  |
|  |  | Scanning Methods : |  |
|  |  | Electronic Convex Sector |  |
|  |  | Electronic Linear |  |
|  |  | Phased |  |
|  |  | Imaging Modes |  |
|  |  | 2D (B-mode), B-mode steering, B-mode AutoAdjust, Colorize 2D, M-mode and PW/CW, PW/CW Doppler, PW/CW Doppler AutoAdjust, C (Color Doppler), TEI (Harmonic Imaging), TPView (Trapezoid Imaging), VPan (Panoramic |  |
|  |  | Imaging), Bidirectional Power Doppler |  |
|  |  | Display Modes |  |
|  |  | 256 gray levels or B-color levels, Orientation: Left/Right, Up/Down, Real-time Triplex mode, 2D+2D (with or without C or PWR D), 2D+M-mode (update or Realtime Duplex), 2D+C+M-mode (update), 2D+Doppler (update or Real-time Triplex), 2D+PWR D, 2D+PWR D+Doppler (update or Real-time Triplex), Colorize on all combinations. |  |
|  |  | Harmonic imaging: should have on all imaging transducers |  |
|  |  | No less than three selectable frequencies for each probe and each scanning mode |  |
|  |  | Speckle reduction and customizable presets for real-time optimization of the image process algorithm |  |
|  |  | Optimize image quality for improve detection of anatomical structures |  |
|  |  | Trapezoid imaging on all linear probes, specially for breast, vascular, musculoskeletal, and thyroid applications |  |
|  |  | Panoramic view |  |
|  |  | Control Panel: Full alphanumeric keyboard, dedicated technology buttons, dedicated function keys. |  |
|  |  | Information displayed on monitor: Application, Selected preset, Date & hour, Type of probe, Probe orientation, Operating frequency range, Acoustic power output, Depth, Focus, Doppler angle, Color & Spectral Doppler filter, Sample volume size, Gain, PW/CW frame rate, Biopsy line, Patient data, Hospital data and annotations Heartbeat, Remote DICOM printing and storage status, Remote digital printing and storage status, Body markers, Timer |  |
|  |  | Image Format: Imaging - Full, Split, Multiple, Left/Right, Up/Down imaging |  |
|  |  | Multi-frequency electronic transducer (Convex, Phased, Linear) |  |
|  |  | In B mode |  |
|  |  | In M mode |  |
|  |  | In PW Doppler mode |  |
|  |  | In CW Doppler mode |  |
|  |  | Measurements and Reporting: all standard in B, M and Doppler modes |  |
|  |  | Abdomen, Breast, Cardiology, Gynecology, Obstetric with programmable tables, Pediatric, Small Parts, Thyroid, Transcranial, MSK, Vascular, Standard biometry reports and user customizable reports, Cardiac, OB/GYN, All reports are automatically stored on the system in the patient file, Archiving of structured reports requires the purchase of the DICOM module |  |
|  |  | Power supply: 220V/50Hz |  |
|  |  | Battery |  |
|  |  | Battery operating (is obligatory): removable Li-ion batteries with operating time not less than 0.5 hour. |  |
|  |  | Probes: 2 probes for abd, brest, ob, gin, small parts: |  |
|  |  | Electronic Convex Sector: from 2 to 5 MHz or equivalent |  |
|  |  | Electronic Linear: from 4.2 to 13 MHz or equivalent |  |
|  |  | Accessories: |  |
|  |  | Ultrasound gel - 1bottle |  |
|  |  | The set includes all the necessary additional devices and accessories to complete functioning of the equipment |  |
|  |  | Standards: |  |
|  |  | ISO13485 or equivalent |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |
| 10 | Անգիոգրաֆիայի համալիր // Floor-mounted digital angiography system for interventional cardiology | General requirements for angiography system | 1 |
|  |  | Floor-mounted single-plane C-arm system for digital imaging techniques, for advanced diagnostic and interventional procedures for angiography, cardiology, electrophysiology and hybrid procedures. |  |
|  |  | High image quality with low dose. |  |
|  |  | Software package including a comprehensive range of applications to enhance image quality without increasing the dose. |  |
|  |  | Flexible configuration capabilities enabling: |  |
|  |  | Interventional cardiology |  |
|  |  | General vascular applications |  |
|  |  | The floormounted supports convenient positioning of the C-arm around the table. |  |
|  |  | C-arm floor base rotation around the table: not less than 200° |  |
|  |  | Patient table a maximum patient weight of not less than 200 kg plus CPR load. The rotating flat detector and collimator shall allow an optimum level of patient access and provide complete anatomical coverage. |  |
|  |  | High resolution, 35 x 35cm, deviation not more than 5cm, solid-state digital flat detector. |  |
|  |  | Should have of combined solutions to reduce exposure of operators and patients, e.g.: |  |
|  |  | reduced pulse rates in fluoroscopy |  |
|  |  | radiation free collimation |  |
|  |  | radiation free patient repositioning |  |
|  |  | automated Cu pre-filtration of X-ray beam |  |
|  |  | display of dose in the examination room |  |
|  |  | Should have of image enhancement tools to increase certainty during interventions, e.g.: |  |
|  |  | image equalization, particularly in areas with large density differences |  |
|  |  | dose adaptive noise reduction |  |
|  |  | vessel edges enhancement |  |
|  |  | detection of fine structures and compensation of motion artifacts |  |
|  |  | image quality customization, according to user’s preference |  |
|  |  | algorithms for improved visibility of stents during cardiac interventions |  |
|  |  | Display of waveforms. Images zoom format. High capacity X-ray tube, enabling higher current during fluoroscopy for difficult and/or obese patients. Angiography system shall allow to control full systems from one workplace (or functionally equivalent feature) and must be configured with high resolution not less four 19" monitors (live/reference/workstation/hemodynamics) medical-grade monitor. |  |
|  |  | The imaging chain allow advanced 2D and 3D applications (possible as option) |  |
|  |  | Real-time high-contrast imaging. |  |
|  |  | Soft tissue imaging with fast reconstruction algorithms |  |
|  |  | Dynamic data update during movements of the C-arm, table, zoom, SID and patient movements compensation. |  |
|  |  | C-arm system ( (not less than presented parametes or analog parametrs) |  |
|  |  | Highly flexible and quick positioning. |  |
|  |  | Integrated patient protection system |  |
|  |  | Isocenter to floor distance: 105 cm, deviation not more than 10% |  |
|  |  | Focus to isocenter distance: 81 cm, deviation not more than 10% |  |
|  |  | Patient coverage: 185cm, deviation not more than 10% |  |
|  |  | C-arm depth: not less than 90 cm |  |
|  |  | Double oblique projections: |  |
|  |  | LAO/RAO: from -120° to +120°, deviation not more than 10% |  |
|  |  | CRAN/CAUD: from + 50° to – 45°, deviation not more than 10% |  |
|  |  | Angulation speed: |  |
|  |  | variable rotation LAO/RAO: not less than 20°/s |  |
|  |  | CRAN/CAUD: not less than 18°/s |  |
|  |  | Variable focal spot to detector distance: from 90cm to 120cm, deviation not more than 10% |  |
|  |  | Patient table |  |
|  |  | Floor-mounted patient table for all angiographic examinations and interventions, large unobstructed cantilevered tabletop and wide range of rotation enabling access to patient from all sides and easy transfer and positioning, table control module for operation of all table functions: |  |
|  |  | Table height adjustable: from 75 cm to 105 cm, deviation not more than 10% |  |
|  |  | Table length: not less than 280cm |  |
|  |  | Table rotation: not less than 240° with increments |  |
|  |  | Longitudinal travel: not less than 110cm |  |
|  |  | Transverse travel: from -15 cm to +15cm , deviation not more than 10% |  |
|  |  | Maximum table load: not less than 200 kg |  |
|  |  | Tabletop: not less than 230 x 45 cm |  |
|  |  | Image system |  |
|  |  | High-resolution digital imaging system with high image quality and real-time image-processing. |  |
|  |  | Image storage capacity: not less than 50,000 images in 1k/12-bit Dmatrix |  |
|  |  | Post processing modes |  |
|  |  | Changing window values. |  |
|  |  | Zooming/Panning: modification on the fly during post processing and pre-configurable for each individual acquisition program. |  |
|  |  | Anatomical background: anatomical surroundings visible by fading in the native image. |  |
|  |  | Electronic shutter: to collimate an image electronically. |  |
|  |  | Annotation: for inserting predefined or free text and drawing lines, arrows and circles. |  |
|  |  | Distance and angle measurement. |  |
|  |  | Pixel shift: manual pixel shift, automatic pixel shift, flexible pixel shift. |  |
|  |  | Operation modes |  |
|  |  | Fluoroscopy |  |
|  |  | Digital pulsed fluoroscopy, with 1.5 or 3.75 to 30p/s in 1k/12-bit matrix (1024x1024 pixels). |  |
|  |  | Last Image Hold |  |
|  |  | Fluoroscopic Loop Imaging |  |
|  |  | Roadmap |  |
|  |  | Individual and serial images in original matrix size, full format and zoom digital real-time filtering, individual; image and series frame rates from 1.5 f/s or 3.75 to 30.0 f/s native, including time-controlled and manually variable frame rates; acquisition, display and storage in original matrix size |  |
|  |  | Digital subtraction angiography in original matrix size, full format and zoom, digital real-time filtering, individual image and series frame rates from 0.5 f/s to 7.5 f/s, including time-controlled and manually variable frame rates; acquisition, display and storage in original matrix size; remask, peak opacification for iodine contrast and CO2 contrast, display of anatomical background from 0 to 100% |  |
|  |  | Cardiac acquisition: acquisition at 7.5, 10, 15 and 30 f/s, acquisition, display and storage in 1k matrix, 12-bit; (including software for enhanced stent visibility). |  |
|  |  | Native 2D-viewing; dynamic subtraction display with optimal alignment of mask and filling and automatic pixel shift over the entire scene. |  |
|  |  | 3D Acquisition: allows native or subtracted 3D for acquisition of 3D high-contrast images. |  |
|  |  | Native 3D reconstruction based on digital rotational angiography with angle triggering. |  |
|  |  | Display and storage ECG and ECG waveform |  |
|  |  | Quantification |  |
|  |  | Vascular analysis for vessel not less diameters of 0.5 mm – 50 mm (not for coronary analysis): measurement program integrated into the imaging system. |  |
|  |  | Left ventricular analysis (must be possible as option) |  |
|  |  | Scientific coronary analysis for vessel diameters: from 0.5 mm to 7 mm, deviation not more than 10% |  |
|  |  | Workplase |  |
|  |  | High-end post processing workstation, for real-time 3D reconstruction and 3D viewing. |  |
|  |  | Full DICOM Functions |  |
|  |  | Archiving function |  |
|  |  | Dose control and monitoring |  |
|  |  | Automatic X-ray control system for fully automatic calculation and optimization of exposure data based on fluoroscopic values |  |
|  |  | Maximum filter equivalent thickness not less than 0.2mm Cu |  |
|  |  | Adjustable Pulsed fluoroscopy |  |
|  |  | Pulse frequency can be adjusted to the requirements of each application to significantly reduce radiation exposure, particularly during interventions |  |
|  |  | Radiation-free positioning of primary and semi-transparent collimators |  |
|  |  | Visually controlled object positioning without radiation; |  |
|  |  | Reference air kerma rate displayed on the data display and image system display. |  |
|  |  | During fluoro pause: accumulated reference air kerma, or dose area product, or percentage of a configurable dose limit value (total of fluoroscopy and acquisition). |  |
|  |  | Low dose acquisition and related software, for dose reduction. |  |
|  |  | DICOM structured dose report. |  |
|  |  | Operation ( (not less than presented function or analog function) |  |
|  |  | In the examination room |  |
|  |  | Complete system operation control at the patient table (C-arm movement, patient table, collimators and software examination protocol). |  |
|  |  | Footswitch for releasing fluoroscopy, acquisition, and table brake |  |
|  |  | In the control room |  |
|  |  | Universal interface using keyboard and mouse for complete system functions such as postprocessing, archiving, and configuring fluoro and acquisition programs. |  |
|  |  | Should have of additional operating options in the control room |  |
|  |  | Flat detector (not less than presented detector or analog detector) |  |
|  |  | Solid-state amorphous silicon 30x30 flat detector or more diagonal entrance plane |  |
|  |  | High-resolution 1k matrix with not more than 200 µm pixel size and not less than 14-bit digitization depth |  |
|  |  | High-speed fiber optic connection to the digital imaging system or equivalent system |  |
|  |  | Integrated temperature stabilizer |  |
|  |  | Integrated collision protection with removable grid |  |
|  |  | Field of view formats: not less than 4 |  |
|  |  | Detector spatial resolution: 2.5 LP/mm, deviation not more than 10% |  |
|  |  | Maximum acquisition speed: 30 f/s , deviation not more than 10% |  |
|  |  | Output digital video matrix: 1024 x 1024, 14-bit , deviation not more than 10% |  |
|  |  | Detector quantum efficiency: not less than 75% |  |
|  |  | X-ray generator (not less than presented generator or analog generator) |  |
|  |  | Microprocessor-controlled high-frequency X-ray generator with automatic dose rate control of fluoroscopy and acquisition |  |
|  |  | Multi-pulse converter frequency: 100 kHz |  |
|  |  | Maximum mA rating: not less than 1000mA, deviation not more than 200mA |  |
|  |  | Tube current range: from 0.5 to 1 mA in 0.01 mA steps , deviation not more than 10% |  |
|  |  | Max. continuous power in fluoro mode: not less than 3000W |  |
|  |  | Tube voltage range for fluoroscopy: from 50 to 120 kV , deviation not more than 10% |  |
|  |  | Automatic X-ray control system for fully automatic calculation and optimization of exposure data based on fluoroscopic values |  |
|  |  | Monitoring of tube load with data display kV and mA post-display on image display |  |
|  |  | Generator control fully integrated in the system control |  |
|  |  | Max. high voltage cable length: 30m , deviation not more than 10% |  |
|  |  | X-ray tube (not less than presented parameters or analog parameters) |  |
|  |  | Not less than two focus high-performance X-ray tube, increased contrast during fluoroscopy, especially for examinations on obese patients: |  |
|  |  | Max. exposure voltage: 125 kV , deviation not more than 10% |  |
|  |  | Focal spot: from 0.4 to 1mm, deviation not more than 0.1mm |  |
|  |  | Anode angle: from 8° to 12° |  |
|  |  | Maximum anode heat content: not less than 3,000,000 HU |  |
|  |  | Maximum heat content of the X-ray tube assembly: not less than 4,900,000 HU |  |
|  |  | Maximum cooling capacity of the anode: not less than 400,000 HU/min. |  |
|  |  | Display ceiling suspension (not less than presented parameters or analog parameters) |  |
|  |  | Ceiling-mounted suspension system for min 4 displays enables height adjustment, longitudinal travel, tilt and swivel capabilities. |  |
|  |  | Length of longitudinal rails: 300 cm , deviation not more than 10% |  |
|  |  | Travel range of ceiling-mounted carriage: < 315 cm |  |
|  |  | Vertical lift (height adjustment): 35 cm , deviation not more than 10% |  |
|  |  | Rotation range of the ceiling-mounted support to the rail axis 300°, deviation not more than 10% settings every 30° , deviation not more than 10% |  |
|  |  | Display (not less than presented parameters or analog parameters) |  |
|  |  | Flat Displays: not less than 3 pcs. |  |
|  |  | TFT high-contrast color display for flicker-free, distortion-free live image and reference image display for X-ray diagnostics as well as interventional therapeutic procedures |  |
|  |  | Light weight, high luminance and contrast values |  |
|  |  | Diagonal screen measurement: 48 cm , deviation not more than 10% |  |
|  |  | Image display: 1280 x 1024 |  |
|  |  | Maximum brightness: 800 cd/m2 , deviation not more than 10% |  |
|  |  | Horizontal viewing area: 170° , deviation not more than 10% |  |
|  |  | Color Display one or more |  |
|  |  | Suitable for color display in the control room |  |
|  |  | Diagonal screen measurement: 48 cm , deviation not more than 10% |  |
|  |  | Image display 1280 x 1024 |  |
|  |  | Maximum luminance: 280 cd/m2 , deviation not more than 10% |  |
|  |  | Viewing angle (min.): 170° H and V , deviation not more than 10% |  |
|  |  | Haemodynamic Monitoring System |  |
|  |  | Bedside and haemodynamic monitoring |  |
|  |  | 7 slot amplifier, one 19' LCD monitor in operation room, two 19' LCD monitors in control room, keyboard, mouse and LaserJet printer |  |
|  |  | IBP: 2 channels, SpO2, NiBP, ECG, ST-analysis |  |
|  |  | Shunts, Resistance, Gradients, valve calculations |  |
|  |  | Diagnostic window |  |
|  |  | Suitable for installation in an on-site wall opening. |  |
|  |  | Frame made of approx. 2 mm thick quality sheet steel |  |
|  |  | Radiation protection glass with lead equivalent: 2.0 mm Pb Lead-glass size (W x H): 120 x 100 cm, deviation not more than 10% |  |
|  |  | Installation data |  |
|  |  | Nominal voltage: (3 ph ± 10%) 380 at 50 Hz ± 1 Hz; |  |
|  |  | Power consumption: 8 kVA for fluoro: 160 kVA for acquisition |  |
|  |  | Accessories: |  |
|  |  | The set includes all the necessary additional devices, accessories to complete functioning of the equipment and safety accessories required for the protection of personnel (example: workstation, contrast injector with 2 syringes, distance remote control and warm function, cover glass, cable, installation materials, insulation materials and ect.) |  |
|  |  | Compatibility to work with angiography scanner |  |
|  |  | Contrast media syringe volume: 150ml, deviation not more than 10% |  |
|  |  | Flow rate range of contrast agent: from 0.1 to 45 ml/sec with step 0.1 ml |  |
|  |  | Step changes in flow rates of contrast agent: from 0.1 to 59.9 ml/min, with step 0,1 ml |  |
|  |  | Color touch screen |  |
|  |  | Auto fill |  |
|  |  | Rise / fall time In a range of from 0 to 9,9 with step 0,1 sec |  |
|  |  | Pressure limit For syringe 150 ml in a range of from 100 to 1200 psi with step 1 psi |  |
|  |  | Fill rate of syringe from 1 to 20 ml/sec |  |
|  |  | Heat maintainer of syringe with contrast agent- Rated temperature 35°С |  |
|  |  | Delay of contrast injection: from 0 to 99,9 with step 0,1 s |  |
|  |  | Duration of injection - Depends on speed and volume |  |
|  |  | Multiphase protocols: 4 phases |  |
|  |  | Ability to perform multiple injections |  |
|  |  | Number of protocols for contrast agent stored in the memory: not less than 40 protocols |  |
|  |  | Electricity power supply 220V, 50Hz |  |
|  |  | Power consumption 1,1 kW |  |
|  |  | Hand Controller to deliver hand injection |  |
|  |  | Functions and properties of patient safety |  |
|  |  | Luminous display function execution protocol |  |
|  |  | Built-in control and regulation of contrast agent pressure |  |
|  |  | Determination of injector head position |  |
|  |  | Ban at the start of injection, at «Syringe up» position of injector head |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |
|  |  | ISO13485 or equivalent |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |

Technical Specifications

Lot 2- Medical Laboratory equipment

| N | NAME | TECHNICAL SPECIFICATION | Q-ty |  |
| --- | --- | --- | --- | --- |
| 1 | Հեմատոլոգիական վերլուծիչ // Hematological analyzer. | Fully automated 5-part differentiation (29 parameters + 3 histograms+3 scattergrams) hematology analyzer | 1 |  |
|  |  | Mode: Whole blood mode, Cappilary whole blood mode, Prediluted mode |  |  |
|  |  | Throughput: not less than 60 samples per hour |  |  |
|  |  | Sample volume: not more than whole blood 15 µl |  |  |
|  |  | User’s interface: not less than 10'' touch screen color TFT display |  |  |
|  |  | Storage for: not less than 40,000 sample results (including histograms) |  |  |
|  |  | Reagent: not more than 3 or 4 reagents |  |  |
|  |  | Flags: Pathological ﬂags (programmable), Lab Limits (programmable) and Reagents alert |  |  |
|  |  | Logs: System Status, Reagents monitoring, Calibration |  |  |
|  |  | Quality control: YES availability of internal and external quality control |  |  |
|  |  | 29 parameters: WBC, Lym#, Mon#, Neu#,Bas#,Eos#, Lym%, Mon%, Neu%,Bas%,Eos%, RBC, HGB, HCT, MCV, MCH, MCHC, RDW-CV, RDW-SD, PLT, MPV, PDW, PCT P-LCR,P-LCC,LIC#,ALY#, LIC%,ALY% and Histogram for WBC, RBC, PLT parameters |  |  |
|  |  | Accessories |  |  |
|  |  | Laser printer - 1pc, keiboard - 1pc, mouse-1pc) |  |  |
|  |  | Sholud have bar code scanner (optional ) |  |  |
|  |  | Start-up Reagents 1000 test |  |  |
|  |  | Power Supply: 220V/50Hz. |  |  |
|  |  | Standards |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 2 | Բիոքիմիական վերլուծիչ // Semi-Automatic Chemistry Analyzer | Semi-Automatic Chemistry Analyzer with incubator | 1 |  |
|  |  | Minimum measurable parameters: ALT, AST, ALP, ACP, GGT, -HBDH, LDH, CK, AMY, TG , TChol, HDL-Chol, UA, UREA, Cr, Glu, TP, Alb, T-bil, D-Bil, TBA, NH4+, HCO3-, Ca, Mg, Cl , Na, K, apoAI , apoB, Lp ( a ), C3, C4 , lgG, lgA, lgM, HbA1c, Alcohol, Amphetamines, Barbiturates, Benzodiazepine |  |  |
|  |  | Analysis method not less than: |  |  |
|  |  | Endpoint, with or without reagent blank |  |  |
|  |  | Endpoint, with sample blank and reagent blank |  |  |
|  |  | Kinetics, with or without reagent blank |  |  |
|  |  | Two-point Kinetics/Fixed-time , with or without |  |  |
|  |  | Reagent blank |  |  |
|  |  | Absorbance |  |  |
|  |  | Turbidimetry |  |  |
|  |  | Linear and non-linear calibration |  |  |
|  |  | Light Source |  |  |
|  |  | Tungsten-halogen Iamp: 6V/10W or 12V/20W |  |  |
|  |  | Wavelength Range |  |  |
|  |  | Automatic by 8-position filter wheel |  |  |
|  |  | Optical filter: minimum 8 wavelengths in the range 340nm to 670nm (with unrepeatable filters) |  |  |
|  |  | Bandwidth: not more than 8nm |  |  |
|  |  | Photometric Range: from 0 to 3.5 ABS |  |  |
|  |  | Stability: not more than 0.005A / 20min at 340nm |  |  |
|  |  | Flowcell |  |  |
|  |  | Stainless steel with quartz window |  |  |
|  |  | Measuring volume: not more than 32 μL |  |  |
|  |  | Optical path: not more than 10mm |  |  |
|  |  | Aspiration volume programmable: from 200 to 9000 μL, deviation not more than 10% |  |  |
|  |  | Touch screen with touch pen and popup keypad |  |  |
|  |  | Mouse or Keyboard via USB |  |  |
|  |  | Printer |  |  |
|  |  | Should have built-in thermal printer |  |  |
|  |  | Should have support with laser jet printers |  |  |
|  |  | Interface |  |  |
|  |  | Should have RS232 serial por |  |  |
|  |  | Should have USB port (for keyboard, mouse or external printer) |  |  |
|  |  | Power Requirement |  |  |
|  |  | Power supply: 220V/50Hz |  |  |
|  |  | Incubator for samples may be internal or external: |  |  |
|  |  | Intenal Incubator: not less than 12 position for 10-15mm round tube |  |  |
|  |  | External Incubator:: not less than 24 position for 10-15mm round tube |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 3 | Կոագուլոմետր // Coagulometer | Coagulometer with: not less than 2 independent operating channels | 1 |  |
|  |  | Semi automatic operation |  |  |
|  |  | Tests performed: PT, APTT, Fibrinogen, cofactors II, V, VII, VIII, IX, X Optical detection with magnetic stirring used to detect clots |  |  |
|  |  | Temperature regulated: 37ºC ± 0.1ºC |  |  |
|  |  | Should have automated test calculation following addition of reagent, automatic compensation for blanks |  |  |
|  |  | Should have straight forward input of test parameters, incubation period, tolerance for test performed in double, reagent data (ISI, unit) |  |  |
|  |  | Should have automatic calibration/self test system |  |  |
|  |  | Capability to store calibration curves, presenting results in different units: second, %, ratio, g/l or mg/ml |  |  |
|  |  | Capacity: PT 60 in duplicate/per hour, APTT 45 in duplicate/per hour |  |  |
|  |  | Should have LCD display |  |  |
|  |  | Incubator capacity: not less than 18 tubes |  |  |
|  |  | RS 232 interface for transferring results to PC or other port for transferring results to PC |  |  |
|  |  | Power supply: 220V/50Hz. |  |  |
|  |  | Accessories: |  |  |
|  |  | Thermo printer - 1pc |  |  |
|  |  | Starter kit consisting of 50 pcs of cuvettes with magnetic stirs |  |  |
|  |  | Consumables: |  |  |
|  |  | PT tests - 1pack |  |  |
|  |  | APTT tests -- 1pack |  |  |
|  |  | Fibrinogen tests -- 1pack |  |  |
|  |  | Printer paper - 5 roll |  |  |
|  |  | Magnetic stirrers - 10 pcs |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 4 | Արյան գազերի և էլեկտրոլիտների վերլուծիչ // Critical Care Analyser | Critical Care Analyser | 1 |  |
|  |  | All sensors are contained in one sensor card or similar technology. |  |  |
|  |  | Blood Gas/Electrolyte/Metabolite: pH, PCO2, PO2, Hct, Na, K, Cl, iCa, Glu, Lac |  |  |
|  |  | Acceptable Samples: Whole blood (heparinized), arterial, mixed venous, capillary |  |  |
|  |  | Sample Volume: Blood Gas/Electrolyte/Metabolite: not more than 100 ul |  |  |
|  |  | Calculated Parameters: SO2%, HCO3-, TCO2, Be-efc, Be-b, SBC, O2Ct, O2Cap, A, AaDO2, a/A, RI, PO2/FIO2, Anion Gap, P50, Hb, nCa |  |  |
|  |  | Temperature Corrected: pH, PCO2, PO2 |  |  |
|  |  | Throughput: not less than 45 Samples/Hour |  |  |
|  |  | Period of service on "board" calibration cartridge not less than 35 days |  |  |
|  |  | Period of storage of the reagents at rom temperature from 18 to 30°C not less than 12 months |  |  |
|  |  | Cartridges: not less than 100 tests |  |  |
|  |  | Accessories: |  |  |
|  |  | Start-up Reagents for 300 tests |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 5 | Կշեռք լաբորատորային // Laboratory scale | Tabletop scales with electronic platforms, scales for recipes, counting scales and also laboratory scales. | 1 |  |
|  |  | Readout: 10 mg |  |  |
|  |  | Measuring range: 600g |  |  |
|  |  | Weighing plate: not less than 130x130 mm or round place not less than 130mm |  |  |
|  |  | Reproducibility: 10 mg, deviation not more than 10% |  |  |
|  |  | Linearity: 30 mg, deviation not more than 10% |  |  |
|  |  | Power supply: 220V/50Hz |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 6 | Ջրային բաղնիք // Water Bath | Capacities: not less than 10 L | 1 |  |
|  |  | Range Ambient: from +5ºC to 100ºC |  |  |
|  |  | Temperature Accuracy: 0.5ºC |  |  |
|  |  | Uniformity: 3.0ºC |  |  |
|  |  | Bath material: seamless stainless steel |  |  |
|  |  | Power Requirements: 220V/50 Hz. |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 7 | Ջրի թորիչ // Water Still | Water still must be removes most dissolved solids and pyrogens from tap water or pretreated water. | 2 |  |
|  |  | Water Still features a thermostat and a water level switch provides automatic operation. |  |  |
|  |  | Protect from over-heating by cut-off when feed water interrupted. |  |  |
|  |  | Easy to clean by removing the condenser from the body. |  |  |
|  |  | The stainless steel double-walled chamber saves electricity. |  |  |
|  |  | Capacity: not less than 10L |  |  |
|  |  | Material: |  |  |
|  |  | Interior material: stainless steel |  |  |
|  |  | Exterior material: stainless steel or steel sheet |  |  |
|  |  | Safety Device: |  |  |
|  |  | Should have water level float switch |  |  |
|  |  | Should have over temp. thermostat |  |  |
|  |  | Should have water supply cut-off valve |  |  |
|  |  | Power source: 220V/50Hz |  |  |
|  |  | Overall (WxDxH): 400x225x545mm |  |  |
|  |  | Standards: |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 8 | Ցենտրիֆուգ // Table Top Centrifuge | Centrifuges are required in the laboratory to separate various components of blood for analysis. | 4 |  |
|  |  | Operational Requirements: |  |  |
|  |  | Aerodynamic compact construction for vibration free performance |  |  |
|  |  | Table top version |  |  |
|  |  | Technical Specifications: |  |  |
|  |  | Tube Capacity: from 5 to 15 ml (tubes vacutainer) have a digital timer |  |  |
|  |  | Body made of strong fabricated & corrosion resistant steel |  |  |
|  |  | Control panel - for start/stop switch, dynamic brakes, step less speed regulator with zero start switch & speed indicator with timer and protective fuses. |  |  |
|  |  | Door interlock |  |  |
|  |  | Maintenance - free brushless drive motor with exact speed pre-selection and display. |  |  |
|  |  | Speed range: from 100 to 5000 rpm, accuracy not less than 1% |  |  |
|  |  | System configuration accessories, spares and consumables |  |  |
|  |  | Centrifuge complete with fixed angle rotor. |  |  |
|  |  | Amount of tube positions in the angle rotor: not less than 8 |  |  |
|  |  | Power Supply: 220V/50Hz |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 9 | Բինոկուլյար մանրադիակ // Microscope Binocular | Quadruple revolving nosepiece. | 4 |  |
|  |  | Optical System: Achromatic optical system |  |  |
|  |  | Binocular Head: Siedentopf Binocular Head, 30°Inclined |  |  |
|  |  | Interpupillary Distance: from 55 to 75 mm, , deviation not more than 10% |  |  |
|  |  | Objectives: Achromatic Objectives 4×; 10×, 40×; 100× |  |  |
|  |  | Eyepiece: Wide Field Eyepiece WF 10X/18 |  |  |
|  |  | Focusing: Coaxial Coarse and Fine Adjustment, Moving Range 24mm, Fine Division 0.002 mm |  |  |
|  |  | Condenser: Abbe Condenser, NA=1.25 with Iris Diaphragm and Filter |  |  |
|  |  | Stage: Double Layers Mechanical Stage X/Y: not less than 140mm×140mm, Moving Range not less than 75×50mm |  |  |
|  |  | Illumination: min. 6V/20W |  |  |
|  |  | Accessories: |  |  |
|  |  | Dust cover, blue and green filter, immersion oil and instruction manual included. 6V 20W halogen bulbs included. |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 10 | Մանրէազերտիչ տաք չոր օդային // Hot air sterilizer | Hot air sterilization | 6 |  |
|  |  | Capacity: not less than 70L |  |  |
|  |  | Temperature Range: from +10 to 200oC, deviation not more than 10% |  |  |
|  |  | Display Resolution: 0.1oC |  |  |
|  |  | Temperature Stability: ±1oC |  |  |
|  |  | Timing Range: from 1 to 999min |  |  |
|  |  | Shelves: not less than 2 |  |  |
|  |  | Should have microprocessor temperature controller |  |  |
|  |  | Should have automatically shut down when the temperature is beyond limit |  |  |
|  |  | Should have air inlet adjustment, timing control, removable shelves and easy chamber cleaning |  |  |
|  |  | Power requirements: 220V/50Hz |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 11 | Թերմոստատ // Incubator | Incubator | 2 |  |
|  |  | Capacity: not less than 50 liters |  |  |
|  |  | Ambient: from +5ºC to 60°C, deviation not more than 10% |  |  |
|  |  | Possibility of continuous running: from 0 to 99 hour |  |  |
|  |  | Should have body made manufactured with stainless steel, insulated with high grade glass wool |  |  |
|  |  | Should have door with gasket and block: |  |  |
|  |  | Should have front see-through inspection window |  |  |
|  |  | Should have heating elements shielded with stainless steel |  |  |
|  |  | Should have adjustable mesh shelves manufactured with stainless steel |  |  |
|  |  | Should have forced air ventilation |  |  |
|  |  | Should have thermometer with analogical dial |  |  |
|  |  | Power Requirements: 220V/50 Hz. |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 12 | Սեղանի ավտոկլավ // Autoclave Table-top | Horizontal autoclave | 2 |  |
|  |  | Self-contained, front-loading and fully automatic |  |  |
|  |  | Manual Door Type |  |  |
|  |  | Material chamber and door: stainless steel |  |  |
|  |  | Capacity: not less than 43L |  |  |
|  |  | Sterilization temperature: not less than 134°C |  |  |
|  |  | Sterilization enhanced by the use of Vacuum Cycle Digital temperature control and timing, Pressure Control , fully programmable cycle parameters |  |  |
|  |  | Low Water alarm |  |  |
|  |  | During Cycle progress: door is locked into position by both temperature and pressure interlocks. |  |  |
|  |  | Electronically-operated Vent valve, Automatic Air Purging, Over Pressure Safety |  |  |
|  |  | Valve and Low Water Cut-Out with Visual and Audible Alarm indication |  |  |
|  |  | Tray: not less than 4 trays |  |  |
|  |  | Tray material: stainless steel |  |  |
|  |  | Power Supply: 220V/50Hz . |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 13 | Ավտոկլավ // Autoclave Steam Sterilizer | Horizontal autoclave | 2 |  |
|  |  | Capacity: not less than 150L |  |  |
|  |  | Sterilization temperature: not less than 134°C |  |  |
|  |  | Self-contained, front-loading unit with integrated steam generator |  |  |
|  |  | Manual Door Type |  |  |
|  |  | Material chamber and door: stainless steel |  |  |
|  |  | During Cycle progress: door is locked into position by both temperature and pressure interlocks |  |  |
|  |  | Should have Automatic Air Purging and Over Pressure Safety |  |  |
|  |  | Power Supply: 3x220V/50Hz . |  |  |
|  |  | Accessories : |  |  |
|  |  | Tray: not less than 4 trays |  |  |
|  |  | Tray material: stainless steel |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 14 | Լաբորատոր գործիքների հավաքածու // Laboratory instrument set | Micropipettors: | 2 |  |
|  |  | Volume: 2-20uL, Increments: 0.02 μL, Error not more than 1% |  |  |
|  |  | Volume: 20-200uL, Increments: 0.2μL, Error not more than 0.8% |  |  |
|  |  | Volume: 100-1000uL, Increments: 1.0μL, Error not more than 0.7% |  |  |
|  |  | Holder for 6 micropipettes - 1pc |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 15 | Իմմունոֆերմենտային հավաքածու (վերլուծիչ, թափահարիչ, լվացող սարք) // ELISA (reader, washer, shaker) | Elisa reader | 1 |  |
|  |  | Micro-plate Reader comes with touch-screen |  |  |
|  |  | Printer (external or internal): availability |  |  |
|  |  | Microwell types: 96 well |  |  |
|  |  | Meserming systems: not less than 8 channel |  |  |
|  |  | Bilt in not less than filters: 405nm, 450nm, 492nm, 630nm, deviation not more than 10% |  |  |
|  |  | Reading range: Automatic focusing function makes the results more precise |  |  |
|  |  | Tray Shaking: availability |  |  |
|  |  | Analysis Mode: Cut-off, single and multi-standard, O.D. bichromatic, dynamic reading or more |  |  |
|  |  | Linear Range: from 0.001 to 3.500Abs |  |  |
|  |  | Connection port: |  |  |
|  |  | RS232 |  |  |
|  |  | USB |  |  |
|  |  | Accessories: Power Cable - 1pc, Fuse - 2pcs, USB Cable - 1pc, Touch Pen - 1pc, Roll of Printer Paper - 2pcs, User Manual - 1pc, Operation Software CD - 1pc |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
|  |  | ELISA Plate Washer |  |  |
|  |  | Residual Volume: not more than 2ul/well |  |  |
|  |  | Dispensing volume: from 50ul to 500ul |  |  |
|  |  | Compatible with 96/48 well plate or strip plate |  |  |
|  |  | Manifolds: 8- or 12- way available |  |  |
|  |  | Wash cycles: from 0 to 99 |  |  |
|  |  | Shaking time: from 0 to 99 seconds |  |  |
|  |  | Alarming: Automatic alarm after washing |  |  |
|  |  | Self-test Function: Automated location check |  |  |
|  |  | Storage: not less than 50 programs |  |  |
|  |  | Interface: RS232 and USB port |  |  |
|  |  | Display: availability |  |  |
|  |  | Accessories: Power Cable - 1pc, Fuse - 2pcs, Tube -1pc, 8-channel manifold - 1pc, Lotion bottle A - 1pc, Waste bottle - 1pc, User Manual - 1pc |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 16 | Մեզի վերլուծիչ // Urinalysis System | Automatic urine colour determination | 1 |  |
|  |  | Built-in printer: availability |  |  |
|  |  | Memory capacity: not less than 100 results |  |  |
|  |  | Interface ports for connection to PC: availability |  |  |
|  |  | Results based upon test strips |  |  |
|  |  | Power Supply: 220V/50Hz. |  |  |
|  |  | Accessories: |  |  |
|  |  | Strips: 100 pcs |  |  |
|  |  | Calibration: 1 box |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |

Technical specifications

Lot 3 - Surgery and Hospital equipment

| N | NAME | TECHNICAL SPECIFICATION REQUIRED BY THE PURCHASER | Q-ty |  |
| --- | --- | --- | --- | --- |
| 1 | ԷՍԳ // Electrocardiograph | Portable | 3 |  |
|  |  | Mains and internal rechargeable battery operation |  |  |
|  |  | Battery operating time: not less than 90 min |  |  |
|  |  | Recording ECG Leads: 12 standard Leads |  |  |
|  |  | Recording Channels: 1/3 user selectable |  |  |
|  |  | LCD: viewing area: 128x 240 mm , deviation not more than 10% |  |  |
|  |  | Lead switching: manual and automatic. |  |  |
|  |  | Sensitivity, mm/mV: minimum 5, 10, 20 |  |  |
|  |  | Calibration signal: automatic end manual. |  |  |
|  |  | Frequency range, Hz Diagnostic: |  |  |
|  |  | Filtered from 0.5 to 30Hz, deviation not more than 10% |  |  |
|  |  | Unfiltered from 0.05 to 150Hz, deviation not more than 10% |  |  |
|  |  | Recorder: INOP control independent for each electrode and pacemaker detection |  |  |
|  |  | Recording method: thermal paper |  |  |
|  |  | Recording paper: 63mm x 30m or other size |  |  |
|  |  | Recording speed, mm/sec: 25/ 50 user selectable |  |  |
|  |  | Channels acquired simultaneously: 1/3 user selectable |  |  |
|  |  | Channels printed simultaneously: 1/3 user selectable |  |  |
|  |  | Power requirements: 220V/50Hz |  |  |
|  |  | Accessories: |  |  |
|  |  | 1. Patient cable |  |  |
|  |  | 2. 6 chest limb electrodes |  |  |
|  |  | 3. 4 limb electrodes |  |  |
|  |  | 4. 4 strap electrodes |  |  |
|  |  | 5. AC/DC adaptor |  |  |
|  |  | 6. 1 bottle ECG Gel |  |  |
|  |  | 7. 2 rolls of paper |  |  |
|  |  | 8. Carry bag |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 2 | Սթրես թեստային համակարգ (տրեդմիլ) // Stress-Test system | Stress ECG system with bilt in electrocardiography module and PC with soft | 1 |  |
|  |  | Built in electrocardiography module |  |  |
|  |  | Sort for tredmell: software for treadmill should be Russified |  |  |
|  |  | Automatic ST measurement; Automatic ECG analysis; Automatic arrhythmia detection and alarm. |  |  |
|  |  | Configuration of analysis parameters; Continuous ECG display; Copies and printed report of analysis result. |  |  |
|  |  | Editing creation of treadmill protocol. |  |  |
|  |  | Easy operation by Touch Screen |  |  |
|  |  | LAN or USB or RS232 connection |  |  |
|  |  | Recording channel: 12/6-ch. |  |  |
|  |  | Recording paper size: A4 |  |  |
|  |  | Recording ECG Leads 12 standard leads: I, II, III, aVR, aVL, aVF, V1, V2, V3, V4, V5, V6 |  |  |
|  |  | Sensitivity: 5, 10, 20mm/mV |  |  |
|  |  | Recording speed: 5, 10, 25, 50mm/sec |  |  |
|  |  | Recording operation: Automatic / Manual |  |  |
|  |  | ECG copy: Reprint for the last patient |  |  |
|  |  | Full disclosure report: Captured ECG trace for 1min while displaying ECG. |  |  |
|  |  | Snap ECG record: Latest signal save for min. 8sec while displaying ECG |  |  |
|  |  | ECG acquisition: Simultaneous 12-lead ECG |  |  |
|  |  | Alphanumeric data input: ID, age, patient name, hospital name and etc with alphanumeric key |  |  |
|  |  | Digital filter: Drift, muscle and AC 50 Hz. |  |  |
|  |  | Calibration: Internal 1 mV signal record for frequency and sensitivity response |  |  |
|  |  | A/D Conversion: not less than 12bits |  |  |
|  |  | Safety Class I (Degree of protection against electric shock: Type CF applied parts) |  |  |
|  |  | Patient Leakage Current: 10 |  |  |
|  |  | A/C Power: 220V/50Hz |  |  |
|  |  | Digital Interface: RS232 to connect external apparatus |  |  |
|  |  | Recorder pressure arterial of blood |  |  |
|  |  | Heart rate: from 30 to 200bpm |  |  |
|  |  | PC network ECG data transmission to PC, ECG save on PC, ECG record to PC printer. |  |  |
|  |  | Real time clock, Date |  |  |
|  |  | Arrhythmia detection: Automatic record mode . |  |  |
|  |  | Pace maker detection. |  |  |
|  |  | Built in Memory: USB/SD card |  |  |
|  |  | Display: LCD not less than 1920x1080 dots, Viewing area: not less than 530x290 mm. |  |  |
|  |  | Computer control treadmill speed and incline |  |  |
|  |  | PC with LCD display, Hard disk with not less than 2GB of free space. |  |  |
|  |  | Treadmill: |  |  |
|  |  | Dimentions: 270x 800x1340mm, deviation not more 10% |  |  |
|  |  | Weight: not more than 220 Kg |  |  |
|  |  | Speed range: from 0.5 Km/h to 20 Km/h , deviation not more 10% |  |  |
|  |  | Accessories: |  |  |
|  |  | PC with LCD display and soft |  |  |
|  |  | Lazer printer for A4 |  |  |
|  |  | Tridmill with built in electrocardiography module |  |  |
|  |  | ECG electrodes - not less than 60pcs |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 3 | Վիրահատական լուսատու առաստաղային (LED տիպի) // OT Light (LED type) | Ceiling mounted system of spring loaded lower arm provides height adjustment for the light head. | 4 |  |
|  |  | Double articulated cardanic system. With horizontal and vertical rotation of the light head ±180°. |  |  |
|  |  | Focusing handle in the sterile area Light head cover easy to clean. |  |  |
|  |  | Aerodynamic shape of the light head reducing disturbance in a laminar flow environment. |  |  |
|  |  | Multi-Colour-chips. Change the colour temperature of the OT-light (colour temperatures not less than 4500 K) |  |  |
|  |  | The surgeon can set the colour temperature according to the tissue structure, the surgical application and individual colour sensitivity. |  |  |
|  |  | Display |  |  |
|  |  | Switching ON and OFF |  |  |
|  |  | Illumination in depth |  |  |
|  |  | Laser pointer |  |  |
|  |  | Electronic light intensity control |  |  |
|  |  | 2 Swivel arms for Axis. |  |  |
|  |  | Combination of two lights on standard axis. Arm 1 endless rotable in all joints, arm 2 with not less than 320o rotation in upper joints, all other joints endless rotable. |  |  |
|  |  | Color temperature: from 3500 to 4750K, adjustable |  |  |
|  |  | Electronic light intensity control at the lamp head: not less than from 5 to 100% |  |  |
|  |  | OT light max. illumination level (1m) – not less than 115,000 lx |  |  |
|  |  | Color Rendering index at 4500K: not less than Ra 96 , R9 not less than 90 |  |  |
|  |  | Focusable size of the light field (in cm): from 17 to 28cm, deviation not more than 10% |  |  |
|  |  | Total power consumption: 70W, deviation not more than 20% |  |  |
|  |  | Number of LEDs: not less than 80 |  |  |
|  |  | Life-span of the LEDs: not less than 40000h |  |  |
|  |  | Working distance (in cm): 60...150cm, deviation not more than 10% |  |  |
|  |  | Diameter of the lamp head (in cm): not less than 49 |  |  |
|  |  | Height adjustment (in cm): not less than 110 |  |  |
|  |  | Satellite: |  |  |
|  |  | OT light max. illumination level (1m) – not less than 115,000 lx |  |  |
|  |  | Color Rendering index at 4500K: not less than Ra 96 , R9 not less than 90 |  |  |
|  |  | Focusable size of the light field (in cm): from 17 to 28cm, deviation not more than 10% |  |  |
|  |  | Total power consumption: 70W, deviation not more than 20% |  |  |
|  |  | Number of LEDs: not less than 80 |  |  |
|  |  | Life-span of the LEDs: not less than 40000h |  |  |
|  |  | Working distance (in cm): from 60 to 150cm, deviation not more than 10% |  |  |
|  |  | Diameter of the lamp head (in cm): not less than 49 |  |  |
|  |  | Height adjustment (in cm): not less than 110 |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 4 | Վիրահատական լուսատու շարժական (LED տիպի) // Mobile OT Light (LED type) | Mobile system of spring loaded lower arm provides height adjustment for the light head. | 1 |  |
|  |  | Double articulated cardanic system. With horizontal and vertical rotation of the light head ±180°. |  |  |
|  |  | Focusing handle in the sterile area. Light head cover easy to clean |  |  |
|  |  | Aerodynamic shape of the light head reducing disturbance in a laminar flow environment. |  |  |
|  |  | Single-Colour-chips. |  |  |
|  |  | Color temperature: approx. 4500K |  |  |
|  |  | Electronic light intensity control at the lamp head: not less than from 5 to 100% |  |  |
|  |  | OT light max. illumination level (1m) – not less than 100,000 lx |  |  |
|  |  | Color Rendering index at 4500K: not less than Ra 95 and R9: not less than 90 |  |  |
|  |  | Focusable size of the light field (in cm): from 17 to 28cm, deviation not more than 10% |  |  |
|  |  | Total power consumption: 30W, deviation not more than 20% |  |  |
|  |  | Number of LEDs: not less than 20 |  |  |
|  |  | Life-span of the LEDs: not less than 40000h |  |  |
|  |  | Working distance (in cm): from 60 to 150cm, deviation not more than 10% |  |  |
|  |  | Diameter of the lamp head (in cm): not less than 49 |  |  |
|  |  | Height adjustment (in cm): not less than 110 |  |  |
|  |  | Power: with integrated power supply 220V |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 5 | Վիրահատական սեղան բազմաֆունկցիոնալ նեռարյալ օրթոպեդիկ կցորդ // Operating table multifunctional with orthopedic attachment | Operating table designed for carrying out treatments, dressing interventions and operations in the range of general surgery and orthopedy. | 1 |  |
|  |  | All metallic parts made from stainless, acid proof steel. |  |  |
|  |  | The table base mobile and have central brakes activated by wire controller. |  |  |
|  |  | Back rest and leg rest inclination angle, Trendelenburg and reverse- |  |  |
|  |  | Trendelenburg positions and height adjustment of the table top activated by electro - hydraulic system |  |  |
|  |  | The table top translucent for x-rays with not less than 5 separate Sections |  |  |
|  |  | Table top length not less than 2130mm |  |  |
|  |  | Table top width not less than 590mm |  |  |
|  |  | Minimal height of the table not less than 720mm |  |  |
|  |  | Maximal height of the table not less than 1080mm indicatively |  |  |
|  |  | Trendelenburg not less than 30º |  |  |
|  |  | Reverse Trendelenburg not less than 30º |  |  |
|  |  | Lateral tilt not less than from -20º to 20º |  |  |
|  |  | Back rest inclination angle not less than from -50º to 70º |  |  |
|  |  | Head rest inclination angle not less than 55º |  |  |
|  |  | Head rest inclination angle not less than 25º |  |  |
|  |  | Battery supply: 12 – 24 V |  |  |
|  |  | Power requirements: 220V/50Hz |  |  |
|  |  | Power Transformer: according to battery voltage |  |  |
|  |  | Accessories: |  |  |
|  |  | Orthopedic attachment (for: Knee arthroscopy; Legs fixation, traction when: Patient in spine position; Patient in lateral position; And for tibia nails; Included transport cart for attachment) |  |  |
|  |  | Anaesthetic Screen with clamp with telescopic tubes: 1pc |  |  |
|  |  | Body restraint Strap with clamp: 1pc |  |  |
|  |  | Padded Shoulder supports: 2pcs |  |  |
|  |  | Padded Leg support with swivel type clamp: 2pcs |  |  |
|  |  | Padded arm Rests 450 -500 mm long with two arm Clamps: 2pcs |  |  |
|  |  | Padded Lateral support with universal attachment Clamp: 2pcs |  |  |
|  |  | Padded rubber mattresses with Anti Microbial agent incorporated into all components that assists in Prohibiting growth of bacteria & fungi and easy to clean and maintain and of 1” thickness: 2X5 Sections, Head rest , I.V drip Stands attachable to the table |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 6 | Վիրահատական սեղան բազմաֆունկցիոնալ // Operating Table Universal | Operating Table Universal (Electric) | 3 |  |
|  |  | All metallic parts made from stainless steel. |  |  |
|  |  | The table base mobile and have central brakes activated by wire controller. |  |  |
|  |  | Table top length not less than 1900mm |  |  |
|  |  | Table top width not less than 500mm |  |  |
|  |  | Minimal height of the table from 700 to 800mm |  |  |
|  |  | Maximal height of the table: from 1000 to 1100mm |  |  |
|  |  | Inclination: |  |  |
|  |  | Should have continuously adjustable back section: from -20o to 45o |  |  |
|  |  | Should have continuously adjustable leg section: from 0o to -90o |  |  |
|  |  | Should have continuously adjustable headrest section: from -30o to 20o |  |  |
|  |  | Power requirements: 220V/50Hz |  |  |
|  |  | Accessories: |  |  |
|  |  | Paper roll holder - 1pc |  |  |
|  |  | Central locking castor system - 1 pc |  |  |
|  |  | Split, removable and spreadable legsetrion - 1pc |  |  |
|  |  | Should have gynecological part - 1pc |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 7 | Ինհալացիոն անզգայացման սարք // Anesthesia trolley (anesthesia machine with ventilator and monitor) | Inhalation anesthesia machine for use with all standard inhalation anesthesia methods, including low-flow. | 4 |  |
|  |  | Precision for pediatric and adults applications. |  |  |
|  |  | Anesthesia machine contained the following parts: Standing compact trolley with minimum three drawers, a pull-out writing tray and front antistatic castors, equipped with brakes; ventilator; manual ventilation unit; breathing system of 3 gases -O2/N2O/AIR- with sensitive oxygen ratio controller for N20 and with the integrated CO2 circle absorber system; integrated VPO monitoring with high resolution color display; interface; socket system for minimum two vaporizers; interface for central gas supply: NIST/ DISS for O2, AIR, N2O included electronical surve and pressure gauges; external fresh gas outlet with all approved non - rebreathing systems, application of all medical gas and volatile gases with manual fresh gas: from 0 to 10 l/min |  |  |
|  |  | Electronic controlled ventilator , electrically driven (no driving gas consumption). |  |  |
|  |  | Operating modes: |  |  |
|  |  | Manual/Spontaneous |  |  |
|  |  | Volume Controlled Ventilation |  |  |
|  |  | Ventilation w/PS (SIMV/PS). |  |  |
|  |  | Pressure Controlled Ventilation; |  |  |
|  |  | Pressure Support; Synchronized Volume Controlled. |  |  |
|  |  | Tidal volume: not less than from 20 to 1400 mL in Volume Control Ventilation and not less than from 20 to 1100mL in SIMV/PS. |  |  |
|  |  | Breathing frequency: not less than 4 to min. 60 bpm |  |  |
|  |  | Max. minute volume (MV): 99 L/min, deviation not more than 5% |  |  |
|  |  | PEEP: from 4 to 30 cmH2O, deviation not more than 5% |  |  |
|  |  | Pressure limiting: from 15 to 70 cmH2O, deviation not more than 5% |  |  |
|  |  | Inspiratory pressure: from 5 to 60 cmH2O, deviation not more than 5% |  |  |
|  |  | Inspiratory pause: from 5 to 50 %, deviation not more than 5% |  |  |
|  |  | Inspiration/expiration ratio: 4 : 1 to 1 : 4 |  |  |
|  |  | SIMV inspiratory time: not less than 0.3 to min. 4.0 sec. |  |  |
|  |  | Trigger: not less than from 0,2 to 10 L/min |  |  |
|  |  | Anesthesia gas delivery system, twin tubes for O2, N2O and Compressed Medical Air. |  |  |
|  |  | Low-flow flowmeter only tube system with: tubes for oxygen from 0 to 0.5L/min and from 0.5 to 12.0 L/min; tubes for nitrous oxide from 0 to 0.5 L/min and from 0.5 to 10.0 L/min; tube for compressed air from 0 to 12 L/min, all deviation not more than 5% |  |  |
|  |  | Have provision for delivery of oxygen, nitrous oxide and medical air with pressure gauges. |  |  |
|  |  | Oxygen and Nitrous oxide linked pneumatically to ensure a minimum of 25% oxygen delivery at all times to avoid delivery of hypoxic mixture. |  |  |
|  |  | Have audio-visual oxygen Failure Warning System with Nitrous oxide cut off. |  |  |
|  |  | Patient monitoring system ECG, Pulseoxymeter and airway pressure, NIBP, rectal/skin temperature. |  |  |
|  |  | Display: color display not less than 10” |  |  |
|  |  | ECG-up to 5 leads, ORS detection range- amplitude 0.5 - 5.0mV; duration - from 40 to 120msec.; pacer detection. |  |  |
|  |  | Respiration measuring range: from 0 to 100 br/min. |  |  |
|  |  | Pulse oximetry measuring range: SpO2 from 1 to 100%; Pulse from -30 to 250 bpm(min). |  |  |
|  |  | Temperature ( absolute and delta temperature): measuring range from 10 to 45°C, deviation not more 0.1°C |  |  |
|  |  | NIBP: Adult from 30 to 240 bpm, deviation not more 5%. Pediatric from 30 to 250mmHg, deviation not more 5%. |  |  |
|  |  | Alarm settings |  |  |
|  |  | Integrated safety functions Sensitive Oxygen Ratio Controller (S-ORC) delivers a minimum O2 concentration of 21% in an O2/N2O mixture. |  |  |
|  |  | N2O cut-off if O2 fresh gas valve is closed or if O2 flow is 0.2 L/min. |  |  |
|  |  | Audible and visual (flashing red LED) indication in case O2 pressure drops below 20 psi (1.38 bar) ± 4 psi (0.27 bar). |  |  |
|  |  | In case of electricity and battery failure, manual ventilation, gas delivery and agent delivery are possible. |  |  |
|  |  | Monitoring Continuous monitoring of inspiratory O2 concentration, breathing frequency, tidal volume, minute volume, peak airway pressure and PEEP, as well as selection of mean or plateau pressure. |  |  |
|  |  | It completed: |  |  |
|  |  | Anesthesia ventilator |  |  |
|  |  | Patient Circuit or patient monitor |  |  |
|  |  | Should have anesthesia gas monitoring system |  |  |
|  |  | Supplied with necessary Attachments for use of the breathing Circuits (Ruben, Bains, Jackson-Rees or Magill) |  |  |
|  |  | Circle Absorber System |  |  |
|  |  | Adjustable pressure limiting Valve, breathing circuit pressure measuring device. |  |  |
|  |  | Gas scavenging and patient suction. |  |  |
|  |  | Bag/Ventilator selecting Valve integrated onto the Absorber. |  |  |
|  |  | Precision, ISO pin type mounting , maintenance free Vaporizers (Temperature, Pressure and Flow compensated) for Izoflurane with standard filling port, with keyed filling device designed for transport with liquid in Vaporiser Chamber with protection against tipping and shaking. |  |  |
|  |  | Power supply 220V/50Hz |  |  |
|  |  | Accessories required for An aesthesia Machine: |  |  |
|  |  | Hose assembly for piped Oxygen supply: 1Set |  |  |
|  |  | Hose assembly for piped Air supply: 1Set |  |  |
|  |  | Test bag, 1 liter: 1 |  |  |
|  |  | Power cord with grounding wire: 1 |  |  |
|  |  | Dust cover: 1 |  |  |
|  |  | Bellows assembly, adult / Children |  |  |
|  |  | Patient Circuit, Complete, Adult / Children. Reusable: Consumables for anesthetic unit for 2 years operation: |  |  |
|  |  | Battery cell (as appropriate) |  |  |
|  |  | Soda lime for circle absorber (5Kg/pack): 5 Packs for 1 unit |  |  |
|  |  | Anesthesia cart - 1pc Designed for using the anesthesia department |  |  |
|  |  | Body in polypropylene |  |  |
|  |  | Should have caster |  |  |
|  |  | Stainless steel |  |  |
|  |  | ABS plastic |  |  |
|  |  | Dimensions: not more than 630x470x1500mm |  |  |
|  |  | The set includes all the necessary additional devices and accessories to complete functioning of the equipment |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 8 | Վիրաբուժական էլեկտրադանակ // Electrosurgical Unit | HF electrosurgical unit used to execute monopolar and bipolar surgery in many fields of application where high precision and reliability are essential | 4 |  |
|  |  | Max. output power for monopolar pure cutting mode - not less than 400W |  |  |
|  |  | Max. output power for monopolar cut coagulation mode - not less than 250W |  |  |
|  |  | Maximum output power monopolar cut mode - not less than 250W |  |  |
|  |  | Max. output for bipolar cutting mode - not less than 100W |  |  |
|  |  | Max. output for bipolar coagulation mode - not less than 100W |  |  |
|  |  | Have thirteen different currents to meet every need : 3 or 4 for monopolar cut, 3 or 4 for monopolar coagulation , 2 for bipolar cut , 3 for bipolar coagulation |  |  |
|  |  | Monopolar cut : |  |  |
|  |  | The currents: pure cut , blend cut with a minimum haemostatic effect and modulated combination of cut and coagulation currents , blend cut with a medium haemostatic effect and modulated combination of cut and coagulation current , blend cut with a strong haemostatic effect spray type and modulated combination of cut and coagulation currents |  |  |
|  |  | Monopolar coagulation: |  |  |
|  |  | Currents, all with automatic optimization system with “under water” capability The four currents: Pin Point Contact , Soft: coagulation , precise, with a minimum superficial effect and deep action, Fulguration: coagulation with superficial effect and deep action. Spray: coagulation without contact, with strong superficial effect and light deep action. Automatic: coagulation with pulsed action |  |  |
|  |  | Bipolar cut |  |  |
|  |  | Two currents: pure cut with a minimum coagulating effect, blend cut with a strong coagulating effect. |  |  |
|  |  | Activation : |  |  |
|  |  | Double pedal switch which used for the monopolar and bipolar functions. Hand-switch handle |  |  |
|  |  | Bipolar electrode with pedal switch or with automatic Start/Stop system (for coagulation only) |  |  |
|  |  | Control |  |  |
|  |  | Control system by microprocessor, monitor continuously all parameters |  |  |
|  |  | Settings : |  |  |
|  |  | Settings and adjustments on the front panel |  |  |
|  |  | Memorization : User able to use at least 4 working programs |  |  |
|  |  | Safety : |  |  |
|  |  | Neutral plate safety circuit , NPCC System : Control connections and contacts of Neutral Plate with Tissues |  |  |
|  |  | Defective Contact notified with visual Alarm and immediate reducing of power |  |  |
|  |  | Generator : |  |  |
|  |  | HF generator with power working according to the principle of free oscillators |  |  |
|  |  | Classification and type: I - CF - leakages on the patient 10μA |  |  |
|  |  | Output circuit: floating - protected against defibrillator interferences. Have HF leakages 150mA |  |  |
|  |  | Power Supply: 220V/50Hz |  |  |
|  |  | Cooling: convection without fan |  |  |
|  |  | Working frequency monopolar and bipolar : from 300 to 600 kHz |  |  |
|  |  | Adjusting powers: by push-buttons |  |  |
|  |  | Accessories connections and activation (with acoustic and luminous signals) |  |  |
|  |  | Accessories : |  |  |
|  |  | Single-use two button Handle |  |  |
|  |  | Autoclavable Handle provided with a 3 Pins socket that fit majority of bipolar electrosurgical units |  |  |
|  |  | Operative Foot-switch (usable as alternative to handle) |  |  |
|  |  | Flat neutral Plate |  |  |
|  |  | Kit of 10 short autoclavable Electrodes 5cm: fine straight, fine angled, loop 4mm, loop 8mm, hook angled, thick angled, ball, ball angled, needle, lancet. |  |  |
|  |  | Bipolar Forceps : 4 |  |  |
|  |  | Forceps provided with a standard European connection |  |  |
|  |  | Straight Forceps – 15 cm. |  |  |
|  |  | Curved Forceps - 15 cm. |  |  |
|  |  | Curved Forceps - 20 cm. |  |  |
|  |  | Straight Forceps - 20 cm. |  |  |
|  |  | Cable, Bipolar |  |  |
|  |  | Adaptor, Bipolar Cable |  |  |
|  |  | User Manual |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 9 | Վիրաբուժական արտածծիչ // Syringe pump | Syringe Driver for Parenteral Infusions, in Wards, OP, ICU's and during Transportation/Transfer | 7 |  |
|  |  | Single Channel. It accept not less than 50cc Syringes |  |  |
|  |  | Non-CaptiveSelf checking of Electronics with audible call back on locked-out Controls |  |  |
|  |  | Tamper and fluid resistant Design. ABS Casing (Impact-Resistant ) |  |  |
|  |  | Flow rate: from 0.1 to 1200 ml/hr, deviation not more than 1% |  |  |
|  |  | Flow rate increment: 0.1 ml/hr |  |  |
|  |  | Total volume infused: from 0.1 to 1000 ml |  |  |
|  |  | Time limit: from 1min. to 99 h. |  |  |
|  |  | Syringes (various brands): from 10 to 50ml |  |  |
|  |  | Precision (stipulated syringes): ± 1% |  |  |
|  |  | Occlusion pressure: adjustable |  |  |
|  |  | KVO: adjustable from 1ml/h to 10ml/h. |  |  |
|  |  | Bolus: adjustable from 0.1 to 1200ml/h |  |  |
|  |  | Alarms/Safety (occlusion; infusion completed; low battery; programming error; syringe positioning error; siring error; clamp error) |  |  |
|  |  | Protection against Leakage Current (type CF. equipment) |  |  |
|  |  | Protection against Electric Shocks (class II, equipment) |  |  |
|  |  | Communication RS232C Interface |  |  |
|  |  | Electrical: |  |  |
|  |  | Electrical class II |  |  |
|  |  | Internal rechargeable battery |  |  |
|  |  | Bat. Autonomy: min. 3 hrs at 5 ml/h |  |  |
|  |  | Power supply: 220V/50Hz |  |  |
|  |  | External transformer. Stabilizer |  |  |
|  |  | General Requirements of Safety for Electromagnetic Compatibility |  |  |
|  |  | Accessories: |  |  |
|  |  | IV stand clamp |  |  |
|  |  | Rail clamp |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 10 | Վիրաբուժական գործիքների հավաքածու // Operation instrument set | Technical description for operation instrument set | 4 |  |
|  |  | NOTE: the following specifications identify the desired configuration of the required operation instrument set and of its components. Bidders are allowed to offer functionally equivalent solutions, including all components described below. |  |  |
|  |  | 1.Holder for gripping and holding the tubular bones, strong-1pc |  |  |
|  |  | 2.Flat chisel with hexagon handle, 15-1pc |  |  |
|  |  | 3.Flat chisel with hexagon handle and with double-sided grinding, 15mm -1pc |  |  |
|  |  | 4.Flat chisel with hexagon handle and with double-sided grinding, 25mm -1pc |  |  |
|  |  | 5. The renal pedicle clamp, curved along the radius N 1, small - 4 pcs. |  |  |
|  |  | 6. Intestinal clip, tight - 2 pcs |  |  |
|  |  | 7. Intestinal flexible clip for adults, curved, 235mm - 2 pcs |  |  |
|  |  | 8. Intestinal flexible clip for adults, straight, 240mm - 2 pcs |  |  |
|  |  | 9. Hemostatic clamp 1x2-toothed,straight N 1 - 12 pcs |  |  |
|  |  | 10. Hemostatic clamp 1x2-toothed,straight N 3 - 4 pcs |  |  |
|  |  | 11. Hemostatic forceps toothed, curved N 1 - 30 pcs |  |  |
|  |  | 12. Hemostatic forceps toothed, straight N 1 - 20 pcs |  |  |
|  |  | 13. Hemostatic MOSQUITO type forceps, curved over a plane -10pcs |  |  |
|  |  | 14. Clip with rack for surgical linen - 12 pcs |  |  |
|  |  | 15. Clip with rack for fixing the linen to the peritoneum - 12 pc |  |  |
|  |  | 16. Bivalve vaginal speculum by Cusco N1 - 1 pc |  |  |
|  |  | 17. Speculum vaginal by Douaen N4 - 1pc |  |  |
|  |  | 18. Speculum vaginal by Douaen N5 - 1pc |  |  |
|  |  | 19. Bilateral speculum by Richardson (from pulmonary surg.) - 2pcs |  |  |
|  |  | 20. Speculum for the abdominal wall with a 100mm width of spoon- 1 pc |  |  |
|  |  | 21. Speculum for the abdominal wall with a 60mm width of spoon- 1 pc |  |  |
|  |  | 22. Speculum for the left lobe of the liver -2 pcs. |  |  |
|  |  | 23. Hepatic speculum 100mm - 1 pc |  |  |
|  |  | 24. Double rector speculum with entire jaws - 1pc |  |  |
|  |  | 25. Rectal speculum for children, 91mm - 1pc |  |  |
|  |  | 26. Probe goitrous with hole - 1pc |  |  |
|  |  | 27. Fluted surgical probe, curved 170 mm- 1 pc |  |  |
|  |  | 28. Fluted surgical probes, bellied (proctologic), 215mm - 1pc |  |  |
|  |  | 29. Bilaterial surgical probes, bellied - 1 pc |  |  |
|  |  | 30. Needle for lumbar puncture A 10х90-30 - 1pc |  |  |
|  |  | 31. Needle for lumbar puncture A 12х120-30 - 1pc |  |  |
|  |  | 32. Blunt ligature big needle , left - 1pc |  |  |
|  |  | 33. Blunt ligature big needle , right - 1pc |  |  |
|  |  | 34. Surgical needle straight with a round hole 0A1-0,7x55 - 50pcs |  |  |
|  |  | 35. Surgical curved needles at 3/8 circle with a triangular tip |  |  |
|  |  | 3В1-0,6х20 -20pcs 3В1-0,7х28 -20pcs 3В1-1,1х50 -20pcs 3В1-1,2х55 -20pcs |  |  |
|  |  | 36. Surgical curved needles at 4/8 circle with a round tip |  |  |
|  |  | 4А1-0,6х20 - 50pcs 4А1-0,6х30 - 50pcs 4А1-0,7х45 - 50pcs 4А1-0,7х50 - 50pcs |  |  |
|  |  | 37. Surgical curved needles at 4/8 circle with a round tip |  |  |
|  |  | 4А1-0,8х32 - 5 pcs. 4А1-0,9х36 - 10 pcs. 4А1-1,0х45 - 10 pcs. |  |  |
|  |  | 38. Surgical curved needles at 4/8 circle with a round tip |  |  |
|  |  | 4А2-0,4х22 - 5pcs |  |  |
|  |  | 39. Surgical curved needles at 4/8 circle with a triangular tip |  |  |
|  |  | 4В1-0,6х20 -30 pcs. |  |  |
|  |  | 4В1-1,0х45 - 20 pcs. |  |  |
|  |  | 4В1-1,2х55 -50 pcs. |  |  |
|  |  | 40. Needle holder, vascular, 220mm - 5 pcs |  |  |
|  |  | 41. Neurosurgical drill with a set of cutters and drills (set) - 1 pc |  |  |
|  |  | 42. Curved forceps - 4pcs |  |  |
|  |  | 43. Straight forceps - 4 pcs |  |  |
|  |  | 44. Sharp tracheotomy hook - 2pcs |  |  |
|  |  | 45. Sharp Surgical hook, 4-toothed Number 2 - 2 pcs |  |  |
|  |  | 46. Blunt Surgical hook, 4-toothed Number 2 - 2 pcs |  |  |
|  |  | 47. Blunt Surgical hook, 4-toothed Number 4 - 2 pcs |  |  |
|  |  | 48. Surgical single-toothed bone hook -2 pcs. |  |  |
|  |  | 49. Surgical hook by Farabeuf (set) - 1 pc |  |  |
|  |  | 50. Cutters by Dahlgren for adults (set) 1 pc |  |  |
|  |  | 51. Bone Cutters for spinal surgery, Liston type with elongated handle 1pc |  |  |
|  |  | 52. Bone Cutters for spinal surgery, Yansen type with elongated handle 1pc |  |  |
|  |  | 53. Two-sided bone spoon, sharp - 2 pcs |  |  |
|  |  | 54. Oval neurosurgical spoons, tough and sharp (set) -1 pc |  |  |
|  |  | 55. Shoulder by Buyalsky for ousting viscera - 1 pcs |  |  |
|  |  | 56. Metal hammer with a rubber grip, medium - 1 pc |  |  |
|  |  | 57. Amputation knife, big НЛ 315x180 - 1 pc |  |  |
|  |  | 58. Amputation knife, small НЛ 250x120 - 1 pc |  |  |
|  |  | 59. Scissors for dissection of the soft tissues in deep civities, vertically-curved, 230mm - 2pcs |  |  |
|  |  | 60. Horizontally curved scissors with a button for cutting bandages - 1 pc |  |  |
|  |  | 61. Vertically curved scissors, blunt, 140 mm - 4 pcs |  |  |
|  |  | 62. Vertically curved scissors, blunt, 170 mm Russian (Нтви- "МИЗ-B") or the same reinforced with hard alloy - 2 pcs |  |  |
|  |  | 63. Surgical blunt scissors, straight 140mm Russian (Нхтп- "МИЗ-B") or the same reinforced with hard alloy - 2 pcs |  |  |
|  |  | 64. Surgical scissors with one sharp tip, 140 mm, Russian (Нхок "МИЗ-B") or the same reinforced with hard alloy - 2 pcs |  |  |
|  |  | 65. Guillotine scissors for ribs - 1 pc |  |  |
|  |  | 66. Twisted Saw Wire - 6 pcs |  |  |
|  |  | 67. Bow-type saw (set) - 1 pc |  |  |
|  |  | 68. General purpose anatomical tweezers ПА 150x2,5 -10 pc |  |  |
|  |  | 69. General purpose anatomical tweezers ПА 200x2,5 -4 pc |  |  |
|  |  | 70. Toothed gripping tweezers ПХ150x5,5 - 2pc |  |  |
|  |  | 71. Ophtalmic surgical forceps , big, straight ПХ100x0,6 - 1pc |  |  |
|  |  | 72. General purpose surgical tweezers ПХ150x2.5 - 10 pcs |  |  |
|  |  | 73. General purpose surgical tweezers ПХ200x2.5 - 4 pcs |  |  |
|  |  | 74. Plate for ousting viscera - 1pcs |  |  |
|  |  | 75. Retractor rack for the chest cavity with of the consumption of the mirrors from 0 to 165 cm. - 1 pc |  |  |
|  |  | 76. Тricuspid retractor with ratchet - 1 pcs |  |  |
|  |  | 77. Curved raspatory, small - 1 pc |  |  |
|  |  | 78. Raspatory, straight - 1 pc |  |  |
|  |  | 79. Ribs raspatories with double bend left - 1 pc |  |  |
|  |  | 80. Ribs raspatories with double bend lright- 1 pc |  |  |
|  |  | 81. Tracheotomy dilator - 1 pcs |  |  |
|  |  | 82. Handles for wire saws (pair) - 1pcs |  |  |
|  |  | 83. Bellied scalpel , medium - Сб150x40 -10pcs |  |  |
|  |  | 84. Sharp tip scalpel, medium Со150x40 - 2pcs |  |  |
|  |  | 85. Scalpel with replaceable blades (1 handle and 10 blades) -2 sets |  |  |
|  |  | 86. Trocars abdominal (set) - 1 pc |  |  |
|  |  | 87.Tracheotomy tubes abdominal №№ 1, 2; 3; 4 ; 5; 6 Titanium for adults or plastic №№ 1, 2; 3; 4 ; 5; 6 - 2 pieces each. |  |  |
|  |  | 88. Neurosurgical spatula two-sided with width of working area 8mm - 1pc |  |  |
|  |  | 89. Neurosurgical spatula two-sided with width of working area 15mm - 1pc |  |  |
|  |  | 90. Forceps for gripping the intestinal wall - 4 pcs |  |  |
|  |  | 91. Intestinal forceps fenestrated for Adults - 1 pc |  |  |
|  |  | 92. Fenestrated hemorrhoidal forceps, straight - 2pcs |  |  |
|  |  | 93. Intestinal forceps, fenestrated for children - 1pc |  |  |
|  |  | 94. Unidental forceps for pulling the uterus (bullet) -1pc |  |  |
|  |  | 95. Sequester forceps, curved N 2 - 1pc |  |  |
|  |  | 96. Bone rongeurs with round gap, curved - 1pc |  |  |
|  |  | 97. Bone rongeurs with round gap, straight - 1pc |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 11 | Թոքերի արհեստական օդափոխության սարք // Ventilator (Intensive ventilation equipment) | Intensive care ventilator with adult, pediatric and neonatal applications. | 3 |  |
|  |  | Color Graphics Display Monitor: not less than 12” color touch monitor |  |  |
|  |  | Not less than present modes: |  |  |
|  |  | Volume Control (Assisted/Controlled) – VCV |  |  |
|  |  | Pressure Control (Assisted/Controlled) – PCV |  |  |
|  |  | Pressure Support – PSV |  |  |
|  |  | Continuous Positive Airway Pressure – CPAP |  |  |
|  |  | SIMV (VCV) + PSV |  |  |
|  |  | SIMV (PCV) + PSV |  |  |
|  |  | Airway Pressure Release Ventilation – APRV |  |  |
|  |  | Pressure Regulated Volume Control – PRVC |  |  |
|  |  | Non-Invasive Ventilation – NIV |  |  |
|  |  | Time Cycled Pressure Limited – TCPL |  |  |
|  |  | SIMV (TCPL) + PSV |  |  |
|  |  | CPAP with Continuous Flow (with leak compensation for NIV) |  |  |
|  |  | Features: |  |  |
|  |  | Inspiratory time: from 0.1 to 25 sec., deviation not more 10% |  |  |
|  |  | Inspiratory flow: from 1 to 180 l/min., deviation not more 10% |  |  |
|  |  | Inspiratory pause: Off, from 0.1 to 2 sec., deviation not more 10% |  |  |
|  |  | Pressure support: from 0 to 60cm H2O/mbar, deviation not more 10% |  |  |
|  |  | Tidal Volume: from 2 to 2000ml., deviation not more 10% |  |  |
|  |  | Resp. rate: from 1 to 100b/min., deviation not more 10% |  |  |
|  |  | Pressure Limit: from 0 to 80cmH2O/mbar., deviation not more 10% |  |  |
|  |  | I: E Rate: inverse: 4:1 or 5:1 |  |  |
|  |  | Trig.(sensitivity) Type: Pressure Trigger from 0 to 20cmH2O/mbar., deviation not more 10% |  |  |
|  |  | Flow Trigger: from 0.1 to 2 l/min., deviation not more 10% |  |  |
|  |  | FiO2: from 0.21 to 1.00 |  |  |
|  |  | PEEP/CPAP: from 0 to 45 cmH2O/mbar ., deviation not more 10% |  |  |
|  |  | Leak Compensation Automatic |  |  |
|  |  | Standby Condition: Allows setting to be preset and Circuit Check tests to be performed prior to starting Ventilation |  |  |
|  |  | Mounted over the Mobile Medical Air Compressor. |  |  |
|  |  | Power requirements: 220 V/50Hz |  |  |
|  |  | Internal Battery provides a minimum 60 minutes for complete ventilator function. |  |  |
|  |  | System Including: |  |  |
|  |  | Medical Air Compressor |  |  |
|  |  | Humidifier |  |  |
|  |  | 1 Tube set, Neonatal |  |  |
|  |  | 1 Tube set, Adult/Children |  |  |
|  |  | 1 Humidifier chamber |  |  |
|  |  | O2-connecting Tube not less than 2 m |  |  |
|  |  | 1 Air-connecting Tube not less than 2 m |  |  |
|  |  | Humidifier: Heating power with safety thermostat and indication of operating status |  |  |
|  |  | 1 Circuit , Neonatal with patient Tubes , water traps, Y-piece, mask elbow, catheter connector |  |  |
|  |  | 1 Circuit , Adult/Children with patient Tubes, water traps, Y-piece, mask elbow, catheter connector |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 12 | Դեֆիբրիլյատոր մոնիտորով // Defibrillator with monitor | Defibrillator with monitor | 4 |  |
|  |  | Portable, lightweight, battery or accumulator operated automated externalto treat patients requiring basic Life Support. Weight of complete unit, excluding batteries or accumulator, permits easy transport. The overall Dimensions enable easy storage and portability. The Casing constructed to withstand the standard operating Conditions in a physician’s office and/or hospital. The device with stand the harshoperating conditions associated with ambulance use. The unit safe to useboth for the operator and the patient. |  |  |
|  |  | Defibrillator |  |  |
|  |  | Operating modes: synchronous and asynchronous, external defibrillation. |  |  |
|  |  | Energy steps: from 1 to 360 joule not less than 22steps |  |  |
|  |  | Discharges: |  |  |
|  |  | 60 shock with 360 joule at 20°C. |  |  |
|  |  | Should have defibrillator Discharge Recovery 5 sec per IEC 60601-2-27 |  |  |
|  |  | Charging time (accumulator fully charged): |  |  |
|  |  | Charging Time to 200J : Within 6 seconds with rated main voltage/DC main Voltage(battery Within 7 seconds ) Charging Time to 360J : Within 8 seconds with rated main voltage/DC main Voltage(battery Within 9 seconds ) |  |  |
|  |  | Delay time: not more than 60 ms between synchronization pulse and energy discharge |  |  |
|  |  | Should have EKG-Monitoring |  |  |
|  |  | ECG monitor: 3 / 5 / 12 Lea |  |  |
|  |  | Lead I, II, III, aVR, aVL, aVF, V1, V2, V3, V4, V5, V6, Paddles, Pads |  |  |
|  |  | Screen Size : not less than 170x128 mm (8.4 in diagonally across the TFT-LCD screen) Screen Type/Color : Liquid Crystal Display (LCD) Color |  |  |
|  |  | Resolution: not less than 800x600 pixel |  |  |
|  |  | Filter: connectable 50Hz |  |  |
|  |  | Functions: |  |  |
|  |  | signal amplification, systole beep, heart rate, paddle lead, energy step, accumulator capacity, heart alarm limits |  |  |
|  |  | Alarm: variable for high and low heart rates |  |  |
|  |  | pulsoximetry – module: |  |  |
|  |  | Indication range: not less than from 100 to 0% |  |  |
|  |  | Calibration range: not less than from 100 to 50% |  |  |
|  |  | Measurement precision: SpO2 |  |  |
|  |  | Adults: not less than from 100 to 70 % ± 2 digits, from 69 to 50 % ± 3 digits |  |  |
|  |  | Operating mode: continuous |  |  |
|  |  | Actualisation time: 2 sec |  |  |
|  |  | Should have printer |  |  |
|  |  | Printer type: thermal transfer head, 1-channel |  |  |
|  |  | Resolution: 8 bit / 200dpi |  |  |
|  |  | Paper width: 80mm |  |  |
|  |  | Operating modes: |  |  |
|  |  | auto print (protocol printing showing the events of the 5s before and 5s after defibrillation) |  |  |
|  |  | ECG print (EKG on-line protocol printing) |  |  |
|  |  | MEMO-print (printing of the saved events of the last 10 defibrillationswith 5s before and 5s after defibrillation) |  |  |
|  |  | 1/3-canal print |  |  |
|  |  | Safety: |  |  |
|  |  | Classification: Protection type II, Type BF, Medical device class 2b |  |  |
|  |  | Power supply: by changeable accumulator |  |  |
|  |  | Changeable accumulator: 14.4V / 6600mAh, Li-ion battery or equivalent accumulator |  |  |
|  |  | Weight: not more than 10kg |  |  |
|  |  | Standards: |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 13 | Ինֆուզոմատ // Volumetric Infusion Pump | Linear peristaltic pump for use with standard from 3.6 to 4.2mm. PVC tubing. | 15 |  |
|  |  | It used equally with infusion sets having standard 20 drops/ml drip chamber or micro-drip - not less than 60 drips/ml drip chamber without the need for flow rate adjustments |  |  |
|  |  | Battery back-up operating time: not less than 4 hours / 125 ml/h |  |  |
|  |  | LCD programming display, data entry calculator style numeric programming keyboard, nurse call output alarm, time and date settings, quick titration of rate or dose with volume-time programming |  |  |
|  |  | Flow rate from 0.1 to 1000ml/h |  |  |
|  |  | Flow rate increment: 0.1 ml/hr |  |  |
|  |  | Volume infused from 1.0 to 9999.9 ml , deviation not more than 1% |  |  |
|  |  | Time limit: from 1min. to 99 h. |  |  |
|  |  | Occlusion pressure: programmable from 0 to 990mm Hg. |  |  |
|  |  | Special functions: titration, fluid balance, clear volume, adjust KVO, adjust bolus |  |  |
|  |  | Programming: ml/h x volume limit; Timex volume limit. |  |  |
|  |  | KVO: programmable/adjustable from 1ml/h to 10ml/h. |  |  |
|  |  | Bolus: programmable/adjustable from 1 to 1000ml/h |  |  |
|  |  | Both flow rates and volume to be infused configured to limit the maximum allowable range accuracy ±5% |  |  |
|  |  | System Configuration Accessories, spares and consumables. |  |  |
|  |  | Power Supply: 220V/50Hz |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 14 | Պացիենտի մոնիտոր // Patient Monitor | Patient monitor draw Cardiograms, measures Blood Pressure not-invasivetype), Respiration values, Temperatures at two spots, Oxygen Saturation inarterial blood (SpO2), and Pulse rate (HR). | 29 |  |
|  |  | Used to Monitor Patients on the transportation by using batteries, and themeasured information on patient's condition displayed in wave forms alongwith numerical values. |  |  |
|  |  | Display/Screen |  |  |
|  |  | Display: Colour TFT display, screen diagonal Curves and measured values: 12”, 8 curves, as well as display fields for Brightness:manually adjusted |  |  |
|  |  | Sweep Speed : adjusted to 6.25, 12.5, 25 and 50 mm/s |  |  |
|  |  | Display pixels: 1024x768 |  |  |
|  |  | Monitor Interfaces |  |  |
|  |  | R-SYNC: for defibrillator, Each R-wave clearly identified, Staff call:Remote alarm via alarm adapter, |  |  |
|  |  | RS-232: as computer interface, Mains/Battery Operation, Mains operation: via built-in Power Supply |  |  |
|  |  | Battery operation: via integrated battery or charging console. |  |  |
|  |  | Battery operating time: 4 hrs |  |  |
|  |  | Power Supply : 220V automatic change-over Power failure |  |  |
|  |  | Buffer: for trend values, |  |  |
|  |  | System Expansion |  |  |
|  |  | Software updates via interfaces: extension of the range of performance parameters |  |  |
|  |  | Monitoring parameters include: ECG, RESP, NIBP, SPO2, PULSE RATE, DUAL-TEMP S-T segment synchronized detection and analysis |  |  |
|  |  | Graphic and tabular trends of all parameters |  |  |
|  |  | NIBP, HEART RATE, TEMP, SPO2, RESP, data storage (400 hour) |  |  |
|  |  | Able to be connected with a Central Monitor Station |  |  |
|  |  | Inner high - definition thermal dot matrix recorder optional, and can output waveforms and characters |  |  |
|  |  | ECG |  |  |
|  |  | Input: Whole-lead ECG cable, standard AAMI cable connector |  |  |
|  |  | Bilt-in printer: availability |  |  |
|  |  | Paper speed: 25, 50mm/s |  |  |
|  |  | Calibration signal: 1mV, |  |  |
|  |  | Protection: Against electrosurgical interference and defibrillation |  |  |
|  |  | Heart Rate: Measuring range: 30…250bpm |  |  |
|  |  | Alarm mode: Audible and visual Alarm |  |  |
|  |  | Measurement Range: 0…99 bpm |  |  |
|  |  | Oxygen Saturation(SPO2) Meas. Range: 35~100% |  |  |
|  |  | NIBP |  |  |
|  |  | Method: Oscillometric, Mode Manual/Auto, Resolution: 1 mmHg |  |  |
|  |  | Have Over-pressure protection: (Adult 300mmHg; Child 220mmHg) Temperature: Means. Range: 0…50oC , Resolution: 0.1oC , Accuracy: ± 0.2oC |  |  |
|  |  | Operation Environment : Temperature: 0~45C , Humidity: 30~85% (non condensing) |  |  |
|  |  | Accessories: |  |  |
|  |  | Electrode cable for disposable electrode (3 leads) |  |  |
|  |  | NIBP cuff for adults |  |  |
|  |  | NIBP cuff for children |  |  |
|  |  | Hose for NIBP cuff |  |  |
|  |  | Power cord |  |  |
|  |  | Disposable electrode (5 pcs) |  |  |
|  |  | SpO2 sensor reusable |  |  |
|  |  | Console / Bedside Holding Device |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 15 | Ավտոմատ ներարկիչ // Syringe pump | Syringe Driver for Parenteral Infusions, in Wards, OP, ICU's and during Transportation/Transfer | 15 |  |
|  |  | Single Channel. It accept not less than 50cc Syringes |  |  |
|  |  | Non-CaptiveSelf checking of Electronics with audible call back on locked-out Controls |  |  |
|  |  | Tamper and fluid resistant Design. ABS Casing (Impact-Resistant ) |  |  |
|  |  | Flow rate: from 0.1 to 1200 ml/hr, deviation not more than 1% |  |  |
|  |  | Flow rate increment: 0.1 ml/hr |  |  |
|  |  | Total volume infused: from 0.1 to 1000 ml |  |  |
|  |  | Time limit: from 1min. to 99 h. |  |  |
|  |  | Syringes (various brands): from 10 to 50ml |  |  |
|  |  | Precision (stipulated syringes): ± 1% |  |  |
|  |  | Occlusion pressure: adjustable |  |  |
|  |  | KVO: adjustable from 1ml/h to 10ml/h. |  |  |
|  |  | Bolus: adjustable from 0.1 to 1200ml/h |  |  |
|  |  | Alarms/Safety (occlusion; infusion completed; low battery; programming error; syringe positioning error; siring error; clamp error) |  |  |
|  |  | Protection against Leakage Current (type CF. equipment) |  |  |
|  |  | Protection against Electric Shocks (class II, equipment) |  |  |
|  |  | Communication RS232C Interface |  |  |
|  |  | Electrical: |  |  |
|  |  | Electrical class II |  |  |
|  |  | Internal rechargeable battery |  |  |
|  |  | Bat. Autonomy: min. 3 hrs at 5 ml/h |  |  |
|  |  | Power supply: 220V/50Hz |  |  |
|  |  | External transformer. Stabilizer |  |  |
|  |  | General Requirements of Safety for Electromagnetic Compatibility |  |  |
|  |  | Accessories: |  |  |
|  |  | IV stand clamp |  |  |
|  |  | Rail clamp |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 16 | Թթվածմնի կոնցենտրատոր //Oxygen Concentrator | Compact and easy to transport oxygen concentrator ( Mobile on Castors ) | 4 |  |
|  |  | The Oxygen Concentrator features a dual-head Compressor. |  |  |
|  |  | One head producing. |  |  |
|  |  | USB Interface |  |  |
|  |  | Volume flow 6 l/min |  |  |
|  |  | O2-concentration: |  |  |
|  |  | from 1 to 4 l/min. not less than 95 % |  |  |
|  |  | from 4 to 5 l/min. not less than85 % |  |  |
|  |  | from 5 to 6 l/min. not less than 75 % |  |  |
|  |  | Flow: |  |  |
|  |  | from 0 to 2 lpm in 0.1 steps |  |  |
|  |  | from 2 to 4 lpm in 0.2 steps |  |  |
|  |  | from 4 to 6 lpm in 0.5 steps |  |  |
|  |  | Pressure-compensated Flow meter permit use of long canula while maintaining accurate flow setting. |  |  |
|  |  | Alarms not less than: temperature, power failure, flow, oxygen, sensor, system |  |  |
|  |  | Pressure-relief: Valve and thermal protection of the Compressor |  |  |
|  |  | Flame-retardant Cabinet |  |  |
|  |  | Sound Level (ANSI): not more than 35dB |  |  |
|  |  | Fixed humidifier Port and Recess prevent bottle and connector breakage |  |  |
|  |  | Power Requirements: 220 V/50 Hz |  |  |
|  |  | Warranty of manufacturer: not less than 30.000 operating hours |  |  |
|  |  | MDD Classification Iia |  |  |
|  |  | Accessories: |  |  |
|  |  | Humidifier, nasal canula: 2 pcs. |  |  |
|  |  | Set of Filters |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 17 | Հոլտեր համակարգ // Holter Recorder with soft and computer | Channels: not less than 12 | 1 |  |
|  |  | Resolution: not less than 10bit |  |  |
|  |  | Recording time: not less than 48hrs |  |  |
|  |  | Recording: Full disclosure |  |  |
|  |  | Record the data on the compact flash card without compression |  |  |
|  |  | Compression: none |  |  |
|  |  | Sampling rate: not less than 128 samples per second |  |  |
|  |  | Frequency response: 0.05Hz to 60Hz -3dB |  |  |
|  |  | Signal verification: Large LCD with lead map |  |  |
|  |  | Pacemaker detection: Programmable ON/OFF with adjustable threshold |  |  |
|  |  | Battery life: availability |  |  |
|  |  | Package Includes: |  |  |
|  |  | Holter monitor (monitoring 24 hours) - 2pcs |  |  |
|  |  | software for ECG recording and USB key - 2pc |  |  |
|  |  | ECG cable with wires 5-lead -2pcs |  |  |
|  |  | Protective case - 2pc |  |  |
|  |  | User guide English - 1pc |  |  |
|  |  | ECG electrodes - not less than 60pcs |  |  |
|  |  | PC with LCD display - 1pc |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 18 | Ճեղքային լամպ // Slit lamp | Professional eye examination slit lamp fully functional; | 1 |  |
|  |  | High Point, comfort is good, the largest 14mm, deviation not less than 10% spot check greater extent of disease |  |  |
|  |  | Professional-grade high-sensitivity lens, Galilean optical design, five-speed drum changefold design more convenient to use |  |  |
|  |  | Should have earn 6X, 10X, 16X, 25X, 40X five kinds of magnification |  |  |
|  |  | Should all lenses through multilayer anti-moisture and mildew, anti-fog coating treatment, in a variety of climatic ring under the environment can ensure a good long-term optical effects |  |  |
|  |  | Should configuration international standard accessory interface, can install various types of posterior segment laser adapter and Goldman tonometer |  |  |
|  |  | Should earmarked digital upgrade interface, you can upgrade to a digital slit lamp |  |  |
|  |  | Eyepiece: 12.5 |  |  |
|  |  | Total magnification: 6x 10x 16x 25x 40x |  |  |
|  |  | Field of View: 6X: 33mm, 10X: 20mm, 16X: 13mm, 25X: 8mm, 40X :5mm |  |  |
|  |  | Range of adjusting the distance of pupil: from 55 mm to 75 mm, deviation not more than 10% |  |  |
|  |  | Range of adjusting eyepieces: from -5D to +5D |  |  |
|  |  | Slit width: from 0 to14mm, deviation not more than 10% |  |  |
|  |  | Slit length: from 1 to14mm, deviation not more than 10% |  |  |
|  |  | Diameter of aperture: φ14mm, φ10mm, φ5mm, φ2mm, φ1mm, φ0.2mm |  |  |
|  |  | Slit angle: from 0 to180 |  |  |
|  |  | Illumination angle: 0, 5, 10, 15, 20 |  |  |
|  |  | Should have filter: Heat absorption, Grey, red free(Green), Cobalt blue |  |  |
|  |  | Bulb: 12v, 50w halogen bulb or LED |  |  |
|  |  | Forward and backward moving: not less than 90mm |  |  |
|  |  | Left and right moving: not less than 100mm |  |  |
|  |  | Handle moving: not less than 15mm |  |  |
|  |  | Up and down moving: not less than 30mm |  |  |
|  |  | Jaw bracket part |  |  |
|  |  | Up and down moving: not less than 80mm |  |  |
|  |  | Should have red led |  |  |
|  |  | Power Supply |  |  |
|  |  | Power input: 220 to 240V/ 50 Hz |  |  |
|  |  | Аccessories |  |  |
|  |  | The set includes all the necessary additional accessories to complete functioning of the equipment (example cover, manual, table with roller for lit lamp and ect) |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 19 | Օֆտալմոսկոպ և օտոսկոպ հավաքածու // Ophthalmoscope and otoscope set | Ophthalmoscope | 1 |  |
|  |  | Ophthalmoscope should have modular design impress with their high-performance optics with aspherical condenser lens and reduced reflection |  |  |
|  |  | Ophthalmoscope should have a thumbrest for securing the ophthalmoscope, which enables easy control of all the elements with one finger. |  |  |
|  |  | Reflexes from the cornea and iris are avoided by separating the observational and illuminating light beams |  |  |
|  |  | Bulb: not less than 2.5 V - 1pc |  |  |
|  |  | Should have 1 spare bulb |  |  |
|  |  | C-handle - 1pc |  |  |
|  |  | Should have 2 alkaline batteries |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
|  |  | Otoscope |  |  |
|  |  | Otoscope with optical glass and reduced refl ection. |  |  |
|  |  | Otoscope with reflection minimized direct illumination not less than 2.5V Xenon. |  |  |
|  |  | Bi-directional swivelling optical glass with 3-fold magnifi cation. |  |  |
|  |  | Suitable for pneumatic tests (supplied without ball) |  |  |
|  |  | C-handle - 1pc |  |  |
|  |  | Should have 2 alkaline batteries |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 20 | Էլեկտրոնեյրոմիոգրաֆ // EMG | Delivery Set | 1 |  |
|  |  | Should have amplifier unit with not less than 2 channels |  |  |
|  |  | Should have holder |  |  |
|  |  | Should have built-in dedicated keyboard (dedicated keyboard for quick access to main commands of EMG software menu items) |  |  |
|  |  | Should have built-in LCD display |  |  |
|  |  | Should have built-in electrical stimulator |  |  |
|  |  | Should have built-in auditory stimulator |  |  |
|  |  | Should have set of EMG electrodes |  |  |
|  |  | Should have surface electrode not less than 2 pcs. |  |  |
|  |  | Should have ground electrode with cable (adult) (400 mm) |  |  |
|  |  | Should have ground electrode with cable (pediatric) (250 mm) |  |  |
|  |  | Should have reusable concentric needle electrode not less than 2 pcs. |  |  |
|  |  | Should have disposable concentric needle electrode not less than 25 pcs |  |  |
|  |  | Should have adapter for needle electrode connection not less than 2 pcs. |  |  |
|  |  | Should have disposable surface electrode not less than 100 pcs |  |  |
|  |  | Should have clip (20 cm) – 2 pcs. (red and black) |  |  |
|  |  | Should have cup electrode with cable – 5 pcs. |  |  |
|  |  | Should have set of stimulators |  |  |
|  |  | Should have measuring reel |  |  |
|  |  | Should have electrode adhesive paste (100 g) |  |  |
|  |  | Should have abrasive paste for skin preparation (160 g) |  |  |
|  |  | Should have software with EMG, EP and Test guide modules for EMG, NCS and SSEP study |  |  |
|  |  | Power Supply |  |  |
|  |  | Power input: 220 to 240V/ 50 Hz |  |  |
|  |  | Аccessories |  |  |
|  |  | The set includes all the necessary additional accessories to complete functioning of the equipment (example computer with monitor, UPS, cover, manual, and ect) |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 21 | ԱՄԲՈՒ // AMBU emergency case | Manual Resuscitator, Autoclavable ,with unidirectional Valve Adapted to First Emergencies the unit and the complement to all automaticventilation | 2 |  |
|  |  | Designed for Adults. Supplied with 2 masks |  |  |
|  |  | All parts autoclavable: up to 134° C. |  |  |
|  |  | The device consist of a self-filling Bag of thin Silicone Rubber whichtransmit adequately the User's Fingertips Action with a high degree of 'feel' indicative of Pressure and Volumes Changes 2 Face masks |  |  |
|  |  | A pressure relief valve incorporated to comply with BS 6850.1987 standart |  |  |
|  |  | Balloon: Silicon volume range not less than 1,6L |  |  |
|  |  | Safety Pressure: not less than 30cm H2O |  |  |
|  |  | Accessories: |  |  |
|  |  | Balloon , autoclavable |  |  |
|  |  | Valve, autoclavable |  |  |
|  |  | Valve Obturator |  |  |
|  |  | PEP Valve Adaptor |  |  |
|  |  | 2 Masks vinyl, Adult / Children with Straps |  |  |
|  |  | Tube for Pressure retake |  |  |
|  |  | Plastic carrying case |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 22 | Օդի ռեցիրկուլյատոր // UV-Air Flow Cleaner–Recirculators Model Recirculator | UV-Air Flow Cleaner–Recirculator: | 23 |  |
|  |  | UV air flow cleaner with 3 UV-lamps |  |  |
|  |  | Full user protection from direct UV radiation |  |  |
|  |  | Convenient fixation on walls (standard) |  |  |
|  |  | Control type: Electronic |  |  |
|  |  | Performance m3/h: 100 |  |  |
|  |  | The bactericidal efficacy: 99.9% |  |  |
|  |  | The total radiation power: 14,4 |  |  |
|  |  | The service life of the UV lamps (h.): 9000 |  |  |
|  |  | Sound power level (dBA): 40 |  |  |
|  |  | Accessories |  |  |
|  |  | Ionized Tube: 3 pcs |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 23 | Նեբուլայզեր // Nebulizer | Ultrasonic nebulizer for aerosol therapy for intensive use. | 2 |  |
|  |  | Ultrasound frequency: from 1.0 to 2.0 MHz, deviation not more than 10% |  |  |
|  |  | Diameter of nebulized particles: from 1 to 5μm, deviation not more than 10% |  |  |
|  |  | Medication cup capacity: 80ml, deviation not more than 10% |  |  |
|  |  | Nebulization rate: 5.0 ml/min at 130 kPa, deviation not more than 10% |  |  |
|  |  | Timer: from 0 to 60min, deviation not more than 10% |  |  |
|  |  | Noise: not more than 50 db |  |  |
|  |  | Power consumption: 220V/50 Hz. |  |  |
|  |  | Electric classification: The product is classified as generic medical Equipment, Class II, Type B |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 24 | Հետազոտման լամպ // Examination Lamp | Lamp, Examination. Mobile on 5 Castors, 5 star-shaped. Heavy and stable stand. Halogen Type | 12 |  |
|  |  | Height of Lamp Head is adjustable |  |  |
|  |  | Light Intensity shall be not less than 50000 Lux at 50sm with beam angle approx. 10o |  |  |
|  |  | Color Intensity: single reflector 3000°K, deviation not more than 10% |  |  |
|  |  | Power rating: min. 50W, 220V/50Hz. |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 25 | Մեծահասակների կծեռք // Scale for Adults | Capacity: from 5 to 150kg | 3 |  |
|  |  | Graduation: 100gr |  |  |
|  |  | Precision: 25gr, deviation not more than 10% |  |  |
|  |  | Height rod: availability |  |  |
|  |  | Washable surface. |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 26 | Լապարասկոպիկ հավաքածու // Laparoscopy set | Technical description for laparoscopic equipment and instruments | 1 |  |
|  |  | NOTE: the following specifications identify the desired configuration of the required laparoscopy system and of its components. Bidders are allowed to offer functionally equivalent solutions, including all components described below. |  |  |
|  |  | 1.Should have digital connection module with the image processing system - 1pc |  |  |
|  |  | 1.1 Resolution pixels: not less than 1920х1080 |  |  |
|  |  | 1.2 Should have progressive scan |  |  |
|  |  | 1.3 Should have video output: DVI-D, 3G - SDI |  |  |
|  |  | 1.4 Connecting additional modules in an amount of not less than 3 |  |  |
|  |  | 1.5 Should have connecting to integrating bus |  |  |
|  |  | 1.6 Should have the presence of a homogeneous lighting mode for endoscopic images |  |  |
|  |  | 1.7 Should have gain mode of contrast for endoscopic images |  |  |
|  |  | 1.8 Number of spectrum shifts of modes for colors of endoscopic images: not less than 2 |  |  |
|  |  | 1.9 Should have withdrawal on one monitor at the same time the two images (Normal + enabled filters) |  |  |
|  |  | 1.10 USB connector for recording pictures, video, and data of the patient in a number not less than 4 |  |  |
|  |  | 1.11 Should have the recording quality is not less than full HD (1080p) |  |  |
|  |  | 1.12 USB flash memory capacity, GB: not less than 32 |  |  |
|  |  | 1.13 Should have safety class is not lower than Cardiac Floating Defib |  |  |
|  |  | 2. Should have digital connection module for 3-chip camera heads - 1pc |  |  |
|  |  | 2.1 Resolution pixels: not less than 1920х1080 |  |  |
|  |  | 2.2 Should have progressive scan |  |  |
|  |  | 2.3 Should have the built-in digital image processing |  |  |
|  |  | 2.4 Should have the presence of a homogeneous lighting mode for endoscopic images |  |  |
|  |  | 2.5 Should have gain mode of contrast for endoscopic images |  |  |
|  |  | 2.6 Number of spectrum shifts of modes for colors of endoscopic images: not less than 2 |  |  |
|  |  | 2.7 Should have safety class is not lower than cardiac floating defib |  |  |
|  |  | 3. Should have 3 chip camera head FULL HD - 1pc |  |  |
|  |  | 3.1 Resolution pixels: not less than 1920x1080 |  |  |
|  |  | 3.2 Should have progressive scan |  |  |
|  |  | 3.3 Should have the built-in digital image processing |  |  |
|  |  | 3.4 Should have the presence of a homogeneous lighting mode for endoscopic images |  |  |
|  |  | 3.5 Should have gain mode of contrast for endoscopic images |  |  |
|  |  | 3.6 Number of spectrum shifts of modes for colors of endoscopic images: not less than 2 |  |  |
|  |  | 3.7 Should have safety class is not lower than cardiac floating defib |  |  |
|  |  | 4. Should have liquid crystal medical flat monitor - 1pc |  |  |
|  |  | 4.1 Should have image format, automatic choosed to the Full HD |  |  |
|  |  | 4.2 Should have sealed metal housing for wet disinfection |  |  |
|  |  | 4.3 Screen Size: not less than 27" |  |  |
|  |  | 4.4 Should have wall mounting with VESA 100 adapter |  |  |
|  |  | 4.5 Aspect ratio: not less than 16:9 |  |  |
|  |  | 4.6 Should have color system PAL / NTSC |  |  |
|  |  | 4.7 Resolution max. pixels: not less than 1920x1080 |  |  |
|  |  | 4.8 Viewing angle of the degrees: not less than 178 |  |  |
|  |  | 4.9 Should have intuitive controls the type of circular slider or other |  |  |
|  |  | 4.10 Should have inputs: 2xDVI, RGB, S-Video, composite video |  |  |
|  |  | 4.11 Should have outputs: DVI, S-Video, composite video |  |  |
|  |  | 4.12 Max. Brightness cd/m2: not less than 240 |  |  |
|  |  | 4.13 Contrast: not less than 1000:1 |  |  |
|  |  | 4.14 Performance for sharp images at movement (ms): not less than 12 |  |  |
|  |  | 4.15 Number of colors / shades bits: not less than 16.8mln |  |  |
|  |  | 4.16 Should have power supply: from 100 to 240 VAC, 50/60 Hz |  |  |
|  |  | 4.17 Power consumption W: not more than 50 |  |  |
|  |  | 4.18 Working temperature °C: from 0 to 40 |  |  |
|  |  | 5. Should have high-power LED source of cold light on LED technology for application in endoscopy in the diagnostic or therapeutic purposes - 1pc |  |  |
|  |  | 5.1 Should have the built-in block of centralized management from a sterile zone for possibility of management from a sterile zone and remote indication of its parameters |  |  |
|  |  | 5.2 Capacity of light (equivalent to xenon) W: 175, deviation not more than 10% |  |  |
|  |  | 5.3 Color temperature of light, K: 6400, deviation not more than 10% |  |  |
|  |  | 5.4 Should have of functions of safety: in case of inclusion the device carries out self-testing |  |  |
|  |  | 5.5 Should have in case of failure detection giving of a sound signal and automatic transition to a safe mode |  |  |
|  |  | 5.6 Should have mandatory ventilation |  |  |
|  |  | 5.7 Diod, which is emitin luminous flux according EN 60825-1 (hermetic frame): Grade 2M |  |  |
|  |  | 5.8 Working temperature °C: from 0 to 40 |  |  |
|  |  | 5.9 Expiration date of a radiator, hours: 30000, deviation not more than 10% |  |  |
|  |  | 5.10 Should have percentage indication of intensity of light |  |  |
|  |  | 5.11 Weght, kg: not more than 5 |  |  |
|  |  | 5.12 Sizes, mm: not more than 305х110х233 |  |  |
|  |  | 5.13 Should have of a waiting mode and memory of the preset value to an exit in a waiting mod |  |  |
|  |  | 6. Fiber-optical light guide for cold light sources - 1pc |  |  |
|  |  | 6.1 Should have the direct plug for connection of the light guide with a light source |  |  |
|  |  | 6.2 Should have especially heat-resistant |  |  |
|  |  | 6.3 Length, cm: 250 |  |  |
|  |  | 6.4 Diameter, mm: 4.8 |  |  |
|  |  | 6.5 Should have sterilization in autoclaves, gas sterilization, chemical disinfection |  |  |
|  |  | 7. Should have coverings protective for use with endoscopic video cameras - 1pc |  |  |
|  |  | 7.1 Should have addition type the telescopic: |  |  |
|  |  | 7.2 Size, cm: not less than 13x242 |  |  |
|  |  | 7.3 Should have sticky tape for fixing |  |  |
|  |  | 7.4 Should have carton insert for insertion |  |  |
|  |  | 7.5 Should have material - transparent polyethylene |  |  |
|  |  | 7.6 Should have disposable, sterile |  |  |
|  |  | 7.7 Should have in packaging, piece: not more than 40 |  |  |
|  |  | 8. Should have antifog solution, ULTRA STOP type, sterile, pierce bottle - 2pcs |  |  |
|  |  | 8.1 Volume ml: 30, deviation not more than 10% |  |  |
|  |  | 9. Should have suction and aspiration roller pump, with laparoscopic and gynecological regimes - 1pc |  |  |
|  |  | 9.1 Should have the built-in block of centralized management from a sterile zone for possibility of management from a sterile zone and remote indication of its parameters |  |  |
|  |  | 9.2 The range of irrigation pressure control in GYN mode, mmHg: from 0 to 200 |  |  |
|  |  | 9.3 The range of irrigation pressure control in LAP mode, mmHg: from 0 to 400 |  |  |
|  |  | 9.4 The range of fluid flow control in GYN mode ml/min: from 0 to 500 |  |  |
|  |  | 9.5 The range of fluid flow control in LAP mode ml/min: from 0 to 1000 |  |  |
|  |  | 9.6 The range of suction pressure control in GYN mode, Bar: from 0 to -0.5 |  |  |
|  |  | 9.7 The range of suction pressure control in LAP mode, Bar: from 0 to -0.8 |  |  |
|  |  | 9.8 Sizes, mm: 310x165x265, deviation not more than 10% |  |  |
|  |  | 9.9 Weght, kg: not more than 6 |  |  |
|  |  | 9.10 Should have automatic recognition of the type of connected tubes and mode selection |  |  |
|  |  | 9.11 Should have conneciont cable to a centralized control system - in complete |  |  |
|  |  | 10. Should have set of silicone suction tube, sterilizable. The set of two pipes. - 1pc |  |  |
|  |  | 10.1 Should have to connect the aspiration not less than 5 liter bottle with pump and suction tool with a bottle |  |  |
|  |  | 10.2 Outer diameter, mm: not nore than 12 |  |  |
|  |  | 10.3 Inner diameter, mm: 8 |  |  |
|  |  | 10.4 The length of the tube from the bottle to the device, cm: 150 |  |  |
|  |  | 10.5 The length of the tube from the bottle to the instrument, cm: 250 |  |  |
|  |  | 11. Should have bottle for suction, for use with endoscopic aspiration-irrigation pumps - 2pcs |  |  |
|  |  | 11.1 Volume, l: not less than 5 |  |  |
|  |  | 11.2 Should have construction material - glass |  |  |
|  |  | 11.3 Should be sterilizable |  |  |
|  |  | 11.4 Should have side scale for monitoring the fill level of aspirate |  |  |
|  |  | 11.5 Line space for monitoring, ml: not less than 250 |  |  |
|  |  | 11.6 Should have used with an airtight lid with a relief valve for overfill protection |  |  |
|  |  | 12. Should have lid for bottle for the endoscopic suction aspiration-irrigation pumps - 2pcs |  |  |
|  |  | 12.1 Should be sterilizable |  |  |
|  |  | 12.2 For the bottles for the aspiration, volume, l: 1.5 or 5 |  |  |
|  |  | 12.3 Should have with safety valve against overflow |  |  |
|  |  | 13. Should have stand for suction bottles -2 pcs |  |  |
|  |  | 13.1 Should have used with glass aspiration bottles |  |  |
|  |  | 13.2 Should have with quickly lockable gate for connecting the holder |  |  |
|  |  | 13.3 Should have material stainless steel |  |  |
|  |  | 13.4 For aspiration bottles, volume, L: not more than 5 |  |  |
|  |  | 14. Should have stand holder for bottles -1pc |  |  |
|  |  | 14.1 Should have for mounting on the metal stand for suction bottle |  |  |
|  |  | 14.2 Should have at the proximal end has a handle |  |  |
|  |  | 14.3 Should have at the distal end has a snap lock |  |  |
|  |  | 14.4 Should have material stainless steel |  |  |
|  |  | 15. Should have foot pedal for irrigation pumps - 1pc |  |  |
|  |  | 15.1 Should have single stage |  |  |
|  |  | 15.2 Should have discrete |  |  |
|  |  | 15.3 Should have single |  |  |
|  |  | 16 Should have bacterial filter for use with the aspiration pumps, disposable, sterile - 3pcs |  |  |
|  |  | 16.1 Should have connection type from the device side - cone |  |  |
|  |  | 16.2 Should have hydrophobic both sides |  |  |
|  |  | 16.3 Should have frame material - polypropylene |  |  |
|  |  | 16.4 Should have filter Material - PTFE |  |  |
|  |  | 16.5 The size of the filter pores, micron: not more than 1 |  |  |
|  |  | 16.6 Packing unit / pack: not less than 10 |  |  |
|  |  | 17. Should have set of irrigation tubes for laparoscopy with two needles for use with roller pumps, disposable - 3pcs |  |  |
|  |  | 17.1 Should have integrated membrane pressure sensor |  |  |
|  |  | 17.2 Should have automatic recognition of the product pump for selecting the operating mode |  |  |
|  |  | 17.3 Length to the irrigation container, mm: 1485, deviation not more than 10% |  |  |
|  |  | 17.4 Length to irrigation instrument, mm: 2545, deviation not more than 10% |  |  |
|  |  | 17.5 inner diameter, mm: not less than 4,7 and not more than 4,9 |  |  |
|  |  | 17.6 Should be sterile |  |  |
|  |  | 17.7 Should have packing, individual, double |  |  |
|  |  | 17.8 Packing unit in package: not less than 10 |  |  |
|  |  | 17.9 Should have latex free |  |  |
|  |  | 17.10 Should have compatible with systems of roller pumps HAMOU ENDOMAT existing to use by the customer |  |  |
|  |  | 18. Should have the high-frequency electrosurgical device, for mono and bipolar modes of coagulation, cutting, vaporization, monopolar and bipolar resectoscopy in normal saline solution - 1pc |  |  |
|  |  | 18.1 Should have with built-in communication block for centralized control from steril area |  |  |
|  |  | 18.2 Power frequency, Hz: not less than 50 - 60 Hz |  |  |
|  |  | 18.3 Min. power consumption, W/VA: 3W/40VA |  |  |
|  |  | 18.4 Min.corrent consumption, mA: 200mA |  |  |
|  |  | 18.5 Max.power consumtion (at 400W): 700W/1150VA |  |  |
|  |  | 18.6 Max.power consumptin (at 400W): 5 A |  |  |
|  |  | 18.7 Input voltage range: 198 V to 264 V |  |  |
|  |  | 18.8 Max.power in the unipolar mode: not less than 400 W |  |  |
|  |  | 18.9 Ma[pPower in bipolar mode: not less than: 400 W |  |  |
|  |  | 18.10 Should have automatic control of the output voltage |  |  |
|  |  | 18.11 Should have mode PWM output voltage |  |  |
|  |  | 18.12 Should have arc density's control mode |  |  |
|  |  | 18.13 Should have digital display control ranges and output parameters on the touch color display |  |  |
|  |  | 18.14 Unipolar sockets: 2x (footswitch and fingerswitch) |  |  |
|  |  | 18.15 Bipolar sockets: 3x (footswitch and fingerswitch |  |  |
|  |  | 18.16 Connection for footswitch: 2 |  |  |
|  |  | 18.17 Should have control regime for neutral electrodes |  |  |
|  |  | 18.18 Should have ability to work with one-surface, double-surface, dynamic, disposable, any neutral electrode |  |  |
|  |  | 18.19 Should have automatic self-test with the issuance of the error code |  |  |
|  |  | 18.20 Should have functions automatic stop / autostart |  |  |
|  |  | 18.21 Should have mode bipolar resection in saline solution |  |  |
|  |  | 18.22 Should have the presence of pre-set parameters of standard procedures |  |  |
|  |  | 18.23 Having the opportunity to introduce and preserve the individual parameters of the user procedures: not less than 300 procedures |  |  |
|  |  | 18.24 Should have bipolar cutting mode: Bit.cutting, Bipolar scissors, Bip.resection, Bip.vaporization |  |  |
|  |  | 18.25 Should have multifunctional bipolar cutting mode: Robi, Laparoscopy, Standard, Bip.resection, Bip.vaporization, BiVascularsafe, BiVascularsafe open surgery, standard auto, Bipolar scissors, Micro, Forsed |  |  |
|  |  | 18.26 Should have unipolar cutting mode: Laparoscopy, Standard, Micro, Resection, Resection C-cut, Supra-Loop, Dry, GastroLoop (1, 2, 3), Gastroknife (1, 2, 3) |  |  |
|  |  | 18.27 Should have unipolar coagulation mode: Laparoscopy, Moderate, Forsed Coag, Resection, Spray, Forsed mixed, Forsed cutting, Gastro coag, Cardiac Thorax, Sim Coag |  |  |
|  |  | 18.28 Compliance with standards: IEC 60601-1: 2005+A1:2012, IEC 60601-1-2:2007, IEC 60601-2-2:2009, ISO 14971:2007, ISO 13485:2003+ Cor.1 2009 |  |  |
|  |  | 18.29 Should have type of protection against electric shock: Class 1 protection |  |  |
|  |  | 18.30 Should have level of protection against electric shock: Type CF |  |  |
|  |  | 18.31 Dimensions in mm: 177х447х457, deviation not more than 10% |  |  |
|  |  | 18.32 Weight : not more than 12,5 kg |  |  |
|  |  | 18.33 Should have included: Power unit, Control cable, Power cord |  |  |
|  |  | 19. Should have threepedal foot switch for the HF generator for switch one mode to another - 1pc |  |  |
|  |  | 20. Should have neutral electrode made of silicone, with 2 rubber strips for fixing, for connection with coagulator used requires connecting cord - 1pc |  |  |
|  |  | 20.1 Contact surface square centimeter: not less than 500 |  |  |
|  |  | 21. Should have connecting cord to connect the neutral electrodes - 1pc |  |  |
|  |  | 21.1 Length, cm: not less than 300 |  |  |
|  |  | 22. Should have monopolar HF cable for devices - 1pc |  |  |
|  |  | 22.1 With plug mm: not less than 5 |  |  |
|  |  | 22.2 Length, cm: not less than 300 |  |  |
|  |  | 23. Should have bipolar cable for resectoscope with a combined highfrequency -mono - and bi connector - 2pcs |  |  |
|  |  | 23.1 With plug mm: not more than 5 |  |  |
|  |  | 23.3 Length, cm: not less than 300 |  |  |
|  |  | 23.4 Should have compatibility with Karl Storz HF electrosurgical device and instruments |  |  |
|  |  | 23.5 Autoclavable into the temperature, Celsium: not less than 134 |  |  |
|  |  | 24. Electronic Thermoflator Set, with integrated communication Module - 1pc |  |  |
|  |  | 24.1 Should have digital indication of modes |  |  |
|  |  | 24.2 Should have gas heating |  |  |
|  |  | 24.3 Range of change of a flow of gas, l/min: from 0 to 30 |  |  |
|  |  | 24.4 Range of adjustment of pressure, mm of: from 0 to 30 |  |  |
|  |  | 24.5 Should have built-in automatic system of safety of fixed control of intra-abdominal pressure, excessive pressure is automatically dumped |  |  |
|  |  | 24.6 Should have visual and acoustic precautionary alarm system |  |  |
|  |  | 24.7 Should have indications of the actual parameters of intra-abdominal pressure, digital |  |  |
|  |  | 24.8 Intra-abdominal pressure of an insufflation in the range, mm of mercury: from 0 to 50mm of mercury |  |  |
|  |  | 24.9 The actual flow of gas, in the range, l/min: from 0 to 30 |  |  |
|  |  | 24.10 The actual consumption of gas in the range, l: from 0 to 999 |  |  |
|  |  | 24.11 Sizes, mm: 305х164х233, deviation not more than 10% |  |  |
|  |  | 24.12 Weight, kg: not more than 7 |  |  |
|  |  | 25. Should have hose of a high pressure with carving connection according to the German standard - 1pc |  |  |
|  |  | 25.1 Length, cm: 55 |  |  |
|  |  | 26. Should have cylinder for CO2, empty, with connection according to the German standard - 1pc |  |  |
|  |  | 27. Should have CO2\N2O the gas filter for use with an insuflator, sterilely - 2pcs |  |  |
|  |  | 27.1 Should have Connection type - ISO |  |  |
|  |  | 27.2 Should have hydrophobic from both sides |  |  |
|  |  | 27.3 Should have frame material - butadiene-sterol copolymer |  |  |
|  |  | 27.4 Should have material of filtre - goffered membrane of glass fiber |  |  |
|  |  | 27.5 Area of filter surfaceа, cm2: not less than 320 |  |  |
|  |  | 27.6 Max speed of gas supply, l/min: not less than 50 |  |  |
|  |  | 27.7 Pack: not less than 25 pcs |  |  |
|  |  | 28. Should have rack mobile for the endoscopic equipment, wide, high - 1pc |  |  |
|  |  | 28.1 Should have on 4 wheels, with brakes |  |  |
|  |  | 28.2 Should have main switch |  |  |
|  |  | 28.3 Number of shelves for equipment: not less than 5 |  |  |
|  |  | 28.4 Number of shelves for equipment, djustable height format: not less than 320 |  |  |
|  |  | 28.5 Should have 2 long side rails for accessories |  |  |
|  |  | 28.6 Should have camera holder |  |  |
|  |  | 28.7 Should have box with the lock |  |  |
|  |  | 28.8 The center stack with integrated power outlets: not less than 12 |  |  |
|  |  | 28.9 The size of the rack, mm: not more than 830 x 1474 x 730 |  |  |
|  |  | 28.10 The size of the shelves, mm: not less than 630 x 25 x 510 |  |  |
|  |  | 28.11 Wheel Diameter, mm: not less than 150 |  |  |
|  |  | 29. Should have the Monitor holder, height-adjustable, bracket VESA 75/100 - 1pc |  |  |
|  |  | 29.1 Rotating degrees: not less than 360 |  |  |
|  |  | 29.2 Load max. kg: not less than 18 |  |  |
|  |  | 30.Should have rigid core - lens optics - 1pc |  |  |
|  |  | 30.1 Should have direct vision, 0 degrees |  |  |
|  |  | 30.2 Diameter, mm: 10 |  |  |
|  |  | 30.3 Should have sapphire face distal lens |  |  |
|  |  | 30.4 Working length, cm: 31 |  |  |
|  |  | 30.5 Should have thermal jack of length |  |  |
|  |  | 30.6 Should have anti-reflex internal covering |  |  |
|  |  | 30.7 Should be autoclavable |  |  |
|  |  | 30.8 Should have with the built-in fiber glass light guide |  |  |
|  |  | 30.9 Large-format, expanded field of vision, corner of the overview of 90 degrees |  |  |
|  |  | 31 Should have rigid core - lens optics |  |  |
|  |  | 31.1 Forward-lateral vision degrees: 30 |  |  |
|  |  | 31.2 Diameter, mm: 10 |  |  |
|  |  | 31.3 Should have sapphire face distal lens |  |  |
|  |  | 31.4 Working length, cm: 31 |  |  |
|  |  | 31.5 Should have thermal jack of length |  |  |
|  |  | 31.6 Should have anti-reflex internal covering |  |  |
|  |  | 31.7 Should be autoclavable |  |  |
|  |  | 31.8 Should have with the built-in fiber glass light guide |  |  |
|  |  | 31.9 Large-format, expanded field of vision, corner of the overview of degrees: 90 |  |  |
|  |  | 32. Should have VERESS needle for pneumoperitoneum - 1pc |  |  |
|  |  | 32.1 Should have the spring mechanism of the stylet |  |  |
|  |  | 32.2 Should have the blunt distal end of the stylet |  |  |
|  |  | 32.3 Should LUER lock |  |  |
|  |  | 32.4 Should be autoclavable |  |  |
|  |  | 32.5 Diameter, mm: 2 |  |  |
|  |  | 32.6 Length, cm: 15 |  |  |
|  |  | 33. Metal trocar, folding on not less 3 parts - 2pc |  |  |
|  |  | 33.1 Should have trocar stylet, pyramidal |  |  |
|  |  | 33.2 Should have cannula without the valve, with the crane for an insufflation |  |  |
|  |  | 33.3 Multipurpose valve, diameter, mm: 11 |  |  |
|  |  | 33.4 Diameter, mm: 11 |  |  |
|  |  | 33.5 Length, cm: 10.5 |  |  |
|  |  | 34. Metal trocar, folding on not less 3 parts - 3pcs |  |  |
|  |  | 34.1 Should have trocar stylet, pyramidal |  |  |
|  |  | 34.2 Should have cannula without the valve, with the crane for an insufflation |  |  |
|  |  | 34.3 Should have multipurpose valve, diameter, mm |  |  |
|  |  | 34.4 Diameter, mm: not more than 6 |  |  |
|  |  | 34.5 Length, cm not more than 10.5 |  |  |
|  |  | 35. Should have trocar with increased valve for insufflation, provides gas flow 30 l/min - 1pc |  |  |
|  |  | 35.1 Should have trocar stylet, pyramidal with oval gas passage |  |  |
|  |  | 35.2 Multipurpose valve, diameter, mm: 11 |  |  |
|  |  | 35.3 Diameter, mm: 11 |  |  |
|  |  | 35.4 Length, cm: 10.5 |  |  |
|  |  | 36. Should have optics Locker, sterile, single use - 1pc |  |  |
|  |  | 36.1 Diameter, mm: 11 |  |  |
|  |  | 36.2 Pack: 12 |  |  |
|  |  | 37. Should have adapter cap with quick fastening to the trocar valve - 1pc |  |  |
|  |  | 37.1 The transition from diameter mm to diameter mm: from 11 to 5 |  |  |
|  |  | 38. Should have suture tool for subcutaneous fascia closing after trocar incisions - 1pc |  |  |
|  |  | 38.1 Diameter, mm: 2.8 |  |  |
|  |  | 38.2 Length, cm: 17 |  |  |
|  |  | 38.3 Should have the handle ring for two fingers |  |  |
|  |  | 39. Should have disassembled Dissecting and Grasping Forceps, rotating - 1pc |  |  |
|  |  | 39.1 Should have both branches are mobile, by KELLY, elongated |  |  |
|  |  | 39.2 Diameter, mm: not more than 5 |  |  |
|  |  | 39.3 Length, cm: 36 |  |  |
|  |  | 39.4 Should have plastic handle, without the racks, with enlarged contact surface of the ring finger |  |  |
|  |  | 39.5 Should have with connection for monopolar coagulation |  |  |
|  |  | 39.6 Should have possibility of fast connection and handle detachment from a tube with an insert by means of a push-button mechanism |  |  |
|  |  | 39.7 Should have the tube with an insert is compatible to all types of handles for monopolar instruments |  |  |
|  |  | 40. Should have forceps for a dissection and coagulation, dismountable - 1pc |  |  |
|  |  | 40.1 Branches bent to the right angle, degrees: 90 |  |  |
|  |  | 40.2 Diameter, mm: not more than 5 |  |  |
|  |  | 40.3 Length, cm: 36 |  |  |
|  |  | 40.4 Should have both branches are mobile |  |  |
|  |  | 40.5 Should have metallic handle, without the racks, with enlarged contact surface of the ring finger |  |  |
|  |  | 40.6 Should have the tube with an insert is compatible to all types of handles for monopolar instruments |  |  |
|  |  | 41. Should have grasping Forceps, rotating , dismountable - 1pc |  |  |
|  |  | 41.1 Diameter, mm: not more than 5 |  |  |
|  |  | 41.2 Length, cm: 36 |  |  |
|  |  | 41.3 Should have plastic handle, with the Manhes type rack, with enlarged contact surface of the ring finger |  |  |
|  |  | 41.4 Should have forceps for a dissection and coagulation, dismountable |  |  |
|  |  | 41.5 Should have branches atraumatic, fenestrated, one is mobile |  |  |
|  |  | 41.6 Should have quick Connect System snap on the tube |  |  |
|  |  | 41.7 Should have the tube with an insert is compatible to all types of handles for monopolar instruments |  |  |
|  |  | 42. Should have forceps for a dissection and coagulation, dismountable - 1pc |  |  |
|  |  | 42.1 Should have branches "alligator" type, both are mobile. |  |  |
|  |  | 42.2 Diameter, mm: not more than 5 |  |  |
|  |  | 42.3 Length, cm: 36 |  |  |
|  |  | 42.4 Should have plastic handle, with the Manhes type rack, with enlarged contact surface of the ring finger |  |  |
|  |  | 43. Should have forceps for a dissection and coagulation, dismountable for working with bowels - 1pc |  |  |
|  |  | 43.1 Should have branches type CROCE-OLMI, atraumatic, fenestrated, slightly curved, one Branche active |  |  |
|  |  | 43.2 Diameter, mm: not more than 5 |  |  |
|  |  | 43.3 Length, cm 36 |  |  |
|  |  | 43.4 Should have plastic handle with disconnectable ratchet, with enlarged contact surface for the finger rings |  |  |
|  |  | 43.5 Should have fast connection with the insertion of the tube with all types of handls with locking snap type |  |  |
|  |  | 44. Should have grasping forcep, dismountable, rotary,with connection for monopolar coagulation - 1pc |  |  |
|  |  | 44.1 Should have fenestrated branches, two active branches for bowel |  |  |
|  |  | 44.2 Should have fast connection with the insertion of the tube with all types of handls with locking snap type |  |  |
|  |  | 44.3 Diameter, mm: not more than 5 |  |  |
|  |  | 44.4 Length, cm: 36 |  |  |
|  |  | 44.5 Should have the tube with an insert is compatible to all types of handles for monopolar instruments |  |  |
|  |  | 44.6 Should have plastic handle, with the Mahnes type rack, with enlarged contact surface of the ring finger |  |  |
|  |  | 45. Should have insert-forceps, to capture the stapler head - 1pc |  |  |
|  |  | 45.1 Diameter, mm: not more than 5 |  |  |
|  |  | 45.2 Length, cm: 36 |  |  |
|  |  | 45.3 Should have fast connection with the insertion of the tube with all types of handls with locking snap type |  |  |
|  |  | 45.4 Should have the tube with an insert is compatible to all types of handles for monopolar instruments |  |  |
|  |  | 46. Should have working insert, biopsy forceps by MAHNES, one branche is active, diameter 5 mm, lenght 36 cm - 1pc |  |  |
|  |  | 47. Should have dismounting dissection and grasping forceps, rotary - 1pc |  |  |
|  |  | 47.1 Diameter, mm: not more than 10 |  |  |
|  |  | 47.2 Length, cm: 36 |  |  |
|  |  | 47.3 Should have branches bent at right angles, the two active jaws |  |  |
|  |  | 47.4 Should have metal handle, without the racks, with enlarged contact surface of the ring finger |  |  |
|  |  | 47.5 Should have with connection for monopolar coagulation |  |  |
|  |  | 47.6 Should have fast connection with the insertion of the tube with all types of handls with locking snap type |  |  |
|  |  | 47.7 Should have the tube with an insert is compatible to all types of handles for monopolar instruments |  |  |
|  |  | 48.Should have working insert, forcers claw-shaped, tooths 2х3 long - 1pc |  |  |
|  |  | 48.1 Diameter, mm: not more than 10 |  |  |
|  |  | 48.2 Length, cm: 36 |  |  |
|  |  | 48.3 Should have the tube with an insert is compatible to all types of handles for monopolar instruments |  |  |
|  |  | 49. Should have forceps, collapsible for BABCOCK, rotary clamp, without connection for monopolar coagulation, two active branches, elongated, reinforced branches - 1pc |  |  |
|  |  | 49.1 Diameter, mm: not more than 10 |  |  |
|  |  | 49.2 Length, cm: 36 |  |  |
|  |  | 49.3 Should have metal handle, without the racks, with enlarged contact surface of the ring finger |  |  |
|  |  | 50. Should have dismountable scissors, rotary, with a connector for monopolar coagulation - 1pc |  |  |
|  |  | 50.1 Should have branche a grooved, curved, conical two mobile branches |  |  |
|  |  | 50.2 Diameter, mm: not more than 5 |  |  |
|  |  | 50.3 Length, cm: 36 |  |  |
|  |  | 50.4 Should have the connection mechanism of the handle and the tube snap |  |  |
|  |  | 50.5 Should have the tube with an insert is compatible to all types of handles for monopolar instruments |  |  |
|  |  | 50.6 Should have plastic handle, without the racks, with enlarged contact surface of the ring finger |  |  |
|  |  | 51. Should have dismountable scissors, rotary, with a connector for monopolar coagulation - 1pc |  |  |
|  |  | 51.1 Should have the blades are bent, two movable branches |  |  |
|  |  | 51.2 Should have the length of the blade mm: 15 |  |  |
|  |  | 51.3 Diameter, mm: not more than 5 |  |  |
|  |  | 51.4 Length, cm: 36 |  |  |
|  |  | 51.5 Should have reusable, autoclavable |  |  |
|  |  | 51.6 Should have button coupling tube to the handle provides a quick and easy assamble and disassemble of instruments |  |  |
|  |  | 51.7 Should have the tube with an insert is compatible to all types of handles for monopolar instruments |  |  |
|  |  | 51.8 Should have plastic handle, without the racks, with enlarged contact surface of the ring finger |  |  |
|  |  | 52. Should have Bipolar Forceps reinforced, rotating, dismountable, model type CLERMONT-FERRAND - 1pc |  |  |
|  |  | 52.1 Should have branche with teeth is especially thin atraumatic, fenestrated with two movable branches |  |  |
|  |  | 52.2 Diameter, mm: not more than 5 |  |  |
|  |  | 52.3 Length, cm: 36 |  |  |
|  |  | 52.4 Should have ring handle, plastic, without fixation |  |  |
|  |  | 52.5 Should have push button mechanism connection handle and the tube insert provides fast assamble and disassemble of instruments |  |  |
|  |  | 52.6 Should have the outher tube with LUER irrigation adapter for cleaning |  |  |
|  |  | 53. Should have bipolar Forceps, rotating, dismountable, model type CLERMONT-FERRAND - 1pc |  |  |
|  |  | 53.1 Should have branches by KELLY slightly curved to the left, with two movablebranches |  |  |
|  |  | 53.2 Diameter, mm: not more than 5 |  |  |
|  |  | 53.3 Length, cm: 36 |  |  |
|  |  | 53.4 Should have push button mechanism connection handle and the tube insert provides fast assamble and disassemble of instruments |  |  |
|  |  | 53.5 Should have ring handle, plastic, without fixation |  |  |
|  |  | 53.6 Should have the outher tube with LUER irrigation adapter for cleaning |  |  |
|  |  | 54. Should have the Dissection and coagulation electrode, L-shaped - 1pc |  |  |
|  |  | 54.1 Diameter, mm: not more than 5 |  |  |
|  |  | 54.2 Length, cm: 36 |  |  |
|  |  | 54.3 Should have tube cm marking |  |  |
|  |  | 55. Should have coagulating and dissection electrode blunt spatula - 1pc |  |  |
|  |  | 55.1 Diameter, mm: not more than 5 |  |  |
|  |  | 55.2 Length, cm: 36 |  |  |
|  |  | 56. Should have the clipapplicator, for use with Ethicon titanic clips LT-300, dismountable,rotated, with a clamp for maintenance branch holding, the size not less than of 10 mm, length of 36 cm, the metal handle with a clamp - 1pc |  |  |
|  |  | 57.Should have hydroperitoneal suction tube, diameter trocar used with 11 mm and 12 mm - 1pc |  |  |
|  |  | 57.1 Diameter, mm: not more than 10 |  |  |
|  |  | 57.2 Length, cm: 36 |  |  |
|  |  | 58. Should have KOH Macro Needle Holder, ergonomic axial handle with disengage able ratchet, ratchet release on the right side - 1pc |  |  |
|  |  | 58.1 Should have straight branches |  |  |
|  |  | 58.2 Diameter, mm: not more than 5 |  |  |
|  |  | 58.3 Length, cm: 33 |  |  |
|  |  | 58.4 Should have carbide-tungsten material of inserts |  |  |
|  |  | 59. Should have KOH Macro Needle Holder, ergonomic axial handle with disengage able ratchet, ratchet release on the right side - 1pc |  |  |
|  |  | 59.1 Diameter, mm: not more than 5 |  |  |
|  |  | 59.2 Length, cm: 33 |  |  |
|  |  | 59.3 Should have carbide-tungsten material of inserts |  |  |
|  |  | 59.4 Should have branches bent to the left |  |  |
|  |  | 60. Should have CICE Knot Tier, model for CLERMONT-FERRAND, for extracorporeal knots - 1pc |  |  |
|  |  | 60.1 Diameter, mm: not more than 5 |  |  |
|  |  | 60.2 Length, cm: 36 |  |  |
|  |  | 61. Should have tube irrigation / aspiration, with anti reflective surface - 1pc |  |  |
|  |  | 61.1 Diameter, mm: not more than 5 |  |  |
|  |  | 61.2 Length, cm: 36 |  |  |
|  |  | 61.3 Should have with the two-way valve for one-hand operation |  |  |
|  |  | 62. Should have endo-loop on ROEDER, with absorbable synthetic suture - 3pcs |  |  |
|  |  | 62.1 USP size: not more than USP 0 |  |  |
|  |  | 62.2 Diameter, mm: not more than 5 |  |  |
|  |  | 62.3 Length, cm: 33 |  |  |
|  |  | 62.4 Should be disposable |  |  |
|  |  | 62.5 Should be sterile |  |  |
|  |  | 62.6 Packing unit / pack: 12 |  |  |
|  |  | 63. Should have gas -filter, disposable with insufflation tube for use with insufflation devices - 2pcs |  |  |
|  |  | 63.1 Should have connection type - ISO |  |  |
|  |  | 63.2 Should have hydrophobic from both sides |  |  |
|  |  | 63.3 Should have patient tube length, m: 3.2 |  |  |
|  |  | 63.4 Should have frame material - butadiene-sterol copolymer |  |  |
|  |  | 63.5 Should have material of filtre - goffered membrane of glass fiber |  |  |
|  |  | 63.6 Should have area of filter surfaceа, cm2: 320 |  |  |
|  |  | 63.7 Max speed of gas supply, l/min: 30 |  |  |
|  |  | 63.8 Pack: 10 |  |  |
|  |  | 63.9 Should have included: an adapter with a pipe, to the non-ISO connection, non-sterile |  |  |
|  |  | 63.10 Tube length, m: 15 |  |  |
|  |  | 63.11 Packing unit / pack: not less than 1 |  |  |
|  |  | 63.12 Should have includes: Hook adapter for devices with a heating element, non-sterile |  |  |
|  |  | 63.13 Packing unit / pack: 1 |  |  |
|  |  | 64. Should have puncture needle, with luer lock - 1pc |  |  |
|  |  | 65. Should have the endo-bag for the extraction of tissues during endoscopic operations - 2pcs |  |  |
|  |  | 65.1 Should have device for putting into the peritoneal cavity with a pusher |  |  |
|  |  | 65.2 Should have included the fishing line Nitinol shape memory in the working hole |  |  |
|  |  | 65.3 Should have self-opening in the abdominal cavity |  |  |
|  |  | 65.4 Should have the two functional sections: Working funnel and tubular container |  |  |
|  |  | 65.5 The diameter of the working section of the disclosure, see: 10 |  |  |
|  |  | 65.6 The total working volume, ml: not less than 260 |  |  |
|  |  | 65.7 The volume of the container section ml: not less than 62 |  |  |
|  |  | 65.8 The length of the container sections, cm: not less than 35 |  |  |
|  |  | 65.9 For use with a trocar size, mm: not less than 10 |  |  |
|  |  | 65.10 Should be sterile |  |  |
|  |  | 65.11 Packing unit / pack: not less than 5 |  |  |
|  |  | 65.12 Should have packing, individual, double |  |  |
|  |  | 65.13 Should have the possibility of using other instruments for the trocar |  |  |
|  |  | 66. Should be uterine manipulator for laparoscopic surgery and pertubation - 1pc |  |  |
|  |  | 66.1 Should have provide anteflektion of uterus and moving towards the inside of the pelvis during laparoscopic gynecologic surgery, including hysterectomy |  |  |
|  |  | 66.2 Uterine anteflektion to degrees: 90 |  |  |
|  |  | 66.3 Should have the ability to lift the uterus to identify the posterior fornix and the posterior vaginal wall |  |  |
|  |  | 66.4 Interchangeable inserts for intrauterine queens size from / to cm: from 5 to 10 |  |  |
|  |  | 66.5 Should have speed marking to determine the position of the uterus |  |  |
|  |  | 66.6 Should have the cervix is retained by a spring retainer for clamping |  |  |
|  |  | 66.7 Should have the locking mechanism should facilitate the manipulation. With the help of the lever arm is fixed and re-released |  |  |
|  |  | 66.8 Should have low weight and ergonomic handle for one-handed operation |  |  |
|  |  | 66.9 Should have connection hose pertubation without sealing expansion and should not allow leak colored liquid |  |  |
|  |  | 66.10 The handle, the weight gr: not more than 400 |  |  |
|  |  | 66.11 Should have working insert 1 |  |  |
|  |  | 66.12 Length mm: not more than 50 |  |  |
|  |  | 66.13 Diameter, mm: not more than 4 |  |  |
|  |  | 66.14 Should have working insert 2 |  |  |
|  |  | 66.15 Length mm: not more than 50 |  |  |
|  |  | 66.16 Diameter, mm: not more than 4,5 |  |  |
|  |  | 66.17 Should have working insert 3 |  |  |
|  |  | 66.18 Length mm: not more than 80 |  |  |
|  |  | 66.16 Diameter, mm: not more than 4,8 |  |  |
|  |  | 66.20 Should have hose holder |  |  |
|  |  | 67. Shoul have pertubation tube for use with the uterine manipulator - 1pc |  |  |
|  |  | 67.1 Should have with the adapter with LUER lock |  |  |
|  |  | 67.2 Should have LUER Lock, female / inner cone for attachment to pipe Ø6mm |  |  |
|  |  | 68. Sould have Uterine Forceps for QUINONES-NEUESTER, with teeth on the branches, with large and small cone with a channel for pertubation - 1pc |  |  |
|  |  | 68.1 Length, cm: 24 |  |  |
|  |  | 69. Should have collapsible exciting forceps rotating - 1pc |  |  |
|  |  | 69.1 Shuld have branche clamping claw, with the teeth 2 x 3 with one movable branche |  |  |
|  |  | 69.2 Diameter, mm: 10 |  |  |
|  |  | 69.3 Length, cm: 36 |  |  |
|  |  | 69.4 Should have Metal handle, with lock type MAHNES |  |  |
|  |  | 69.5 Should have snap handle locking mechanism with tube insert |  |  |
|  |  | 70. Should have plastic container for storage and sterilization - 1pc |  |  |
|  |  | 70. 1 Should have transparent cover with lock to lock |  |  |
|  |  | 70.2 Removable stand and paste tools to store up to 12 instruments with diameters from 2.5 to 10 mm and up to 6 trocars |  |  |
|  |  | 70.3 Should have perforation pad |  |  |
|  |  | 70.4 External dimensions mm: not more than 532 x 254 x 165 |  |  |
|  |  | 71. Should have retainer for fibroids, for use with a trocar 6 mm - 1pc |  |  |
|  |  | 71.1 Should ave with corkscrew tip |  |  |
|  |  | 71.2 Should have with extended tip |  |  |
|  |  | 71.3 Diameter, mm: not more than 5 |  |  |
|  |  | 71.4 Length, cm: not more than 36 |  |  |
|  |  | 72. Should have basket-screen average, for the cleaning, sterilization and storage of the two endoscopes and one fiber optic cable - 1pc |  |  |
|  |  | 72.1 Sizes, mm: 487х125х54, deviation not more than 10% |  |  |
|  |  | 72.2 For rigid endoscopes of diameter up to mm: 10 |  |  |
|  |  | 72.3 And the working length to cm: 32 |  |  |
|  |  | 73. Should have oiler - drip - 1pc |  |  |
|  |  | 73.1 Should have it has the form of a cylinder equipped with a valve control button in the proximal portion and the spout in the form of a tube in the distal part |  |  |
|  |  | 73.2 Should have lockable threaded cover for closing the drip spout |  |  |
|  |  | 74. Should have oil for lubrication of instruments in the container - 1pc |  |  |
|  |  | 74.1 Volume, ml: 50 |  |  |
|  |  | 74.2 Should have possibility of to autoclave lubricated instruments |  |  |
|  |  | 75. Should have special grease for taps - 1pc |  |  |
|  |  | 75.1 Should have the possibility of using endoscopes for the lubrication of cranes |  |  |
|  |  | 75.2 Weght, gr: 3 |  |  |
|  |  | 76. Should have cleaning brush for sharp branches - 1pc |  |  |
|  |  | 76.1 Should have formed as a plate having a distal portion of the bristles on one side directed perpendicular to the longitudinal axis of the main plane of the plate |  |  |
|  |  | 76.2 The size of the working part, cm: 4 |  |  |
|  |  | 76.3 The height of the bristles, cm: not more than 1,7 |  |  |
|  |  | 77. Should have cleanser for distal and proximal lenses and fiber optic endoscope surfaces, a tube - 1pc |  |  |
|  |  | 77.1 Weght, gr: 5 |  |  |
|  |  | 78. Should have cleaning pistol with cleaning attachments included for quick and easy washing or drying instruments - 1pc |  |  |
|  |  | 78.1 Should have with flexible armored hose |  |  |
|  |  | 78.2 Should have connected to the tap water or air supply through a standard 3/4 "jack |  |  |
|  |  | 78.3 Max. air or water pressure, bar: not more than 6 |  |  |
|  |  | 78.4 Should have pistol with thermal protection |  |  |
|  |  | 78.5 Should be autoclavable |  |  |
|  |  | 78.6 Should have wall tool holder |  |  |
|  |  | 78.7 Number of nozzles for washing various types of instruments, pieces: not more than 8 |  |  |
|  |  | 78.8 Should have the nozzle for cleaning the gun for cleaning the cannula |  |  |
|  |  | 78.9 Should have nozzle cleaning gun cleaning instrumentsl channels |  |  |
|  |  | 78.10 Should have the nozzle for cleaning the gun to clean the valves |  |  |
|  |  | 78.11 Should have nozzle for a cleaning pistol for cleaning silicone tubing |  |  |
|  |  | 78.12 Should have the nozzle for cleaning the gun for cleaning LUER cones |  |  |
|  |  | 78.13 Should have nozzle for a cleaning pistol for cleaning surfaces and cavities |  |  |
|  |  | 78.14 Should have nozzle for a cleaning pistol for cleaning surfaces and internal cavities |  |  |
|  |  | 78.15 Should have nozzle for the cleaning pistol as the elevator fluid aspiration from the corners and cavities |  |  |
|  |  | 79. Should have nozzle for the cleaning pistol as the elevator liquid cleaning brush, tube - 20pcs |  |  |
|  |  | 79.1 Length, cm: 50 |  |  |
|  |  | 79.2 Diameter, mm: not more than 2.5 |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 27 | Ճակատային լամպ վիրաբուժական // Head ligth | Glass fiber cable for cold-light supply with centered cable duct for undisturbed working. | 1 |  |
|  |  | Ergonomic headband with head bow and removable leather padding for comfortable use even over several hours, infinite height and width adjustment with central wheels adjustable with one hand. |  |  |
|  |  | Should have projector |  |  |
|  |  | Projector light output: not less than 100W xenon |  |  |
|  |  | Should have 1 spare bulb |  |  |
|  |  | Fast push connector for connecting glass fiber cable and light source. |  |  |
|  |  | Light exit: focusable, double lens system, ball joint. |  |  |
|  |  | Glass fiber cable: |  |  |
|  |  | Length not less than 220 cm |  |  |
|  |  | Diameter not less than 3.5 mm |  |  |
|  |  | Accessories: |  |  |
|  |  | The set includes all the necessary additional devices, accessories to complete functioning of the equipment |  |  |
|  |  | Powerful head light with focusable LED lamp. |  |  |
|  |  | Headband version, with built-in battery compartment and integral charge jack, can be optionally fi tted with 2 lithium batteries size CR 123 A, 4 batteries size AAA (alkaline) or 4 rechargeable batteries size AAA (NiMH). |  |  |
|  |  | Should have particularly comfortable with balanced, infinitely adjustable headband. |  |  |
|  |  | Should be use without any hindering cables. |  |  |
|  |  | Should have On/off switch on battery compartment. |  |  |
|  |  | Charge jack for plug-in charger integrated in battery compartment. |  |  |
|  |  | White LED: not more than 6 V/ 1 W, not less than 6.200 Kelvin |  |  |
|  |  | Battery operation: not less than 3 hrs. with lithium batteries, not less than 90 min. with alkaline batteries Type AAA. |  |  |
|  |  | Rechargeable battery operation: not less than 90 min. with fully charged battery. |  |  |
|  |  | Accessories: |  |  |
|  |  | The set includes all the necessary additional devices, accessories to complete functioning of the equipment |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 28 | Նեյրովիրաբուժական մեծ հավաքածու // Neurosurgical Instrument set | NOTE: the following specifications identify the desired configuration of the required neurosurgical instrument set and of its components. Bidders are allowed to offer functionally equivalent solutions, including all components described below. | 1 |  |
|  |  | 1. Neurosurgical hemostatic clamp, straight -10 pcs |  |  |
|  |  | 2. Neurosurgical hemostatic clamp, curved- 30 pcs |  |  |
|  |  | 3. Hemostatic MOSQUITO type clamp, straight - 6 pcs |  |  |
|  |  | 4. Hemostatic MOSQUITO type clamp curved over a plane - 6pcs |  |  |
|  |  | 5. Hemostatic MOSQUITO type clamp curved along the edge/rib-6pcs |  |  |
|  |  | 6. Valleculate surgical probe, length 170 mm - 2pcs |  |  |
|  |  | 7. Blunt ligature needle , right N3 - 1pc |  |  |
|  |  | 8. Blunt ligature big needle , left N3- 1pc |  |  |
|  |  | 9. Bifurcated ligature needle - 1 pc |  |  |
|  |  | 10. Surgical curved needles at 4/8 circle with a round hole 4А1-0,6х30 -40pcs 4А1-0,6х20 -40pcs |  |  |
|  |  | 12. Surgical curved needles at 4/8 circle with a triangular tip 4B1-0,7х28 -40pc 4B1-0,6х20 -40pc |  |  |
|  |  | 13. Needle for lumbar puncture A-1,0х90 30 - 10pc |  |  |
|  |  | 14. Needle for lumbar puncture A-1,2х120 30 - 10pc |  |  |
|  |  | 15. Rongeurs by Jegoroff-Freidin N1 -1 pc |  |  |
|  |  | 16. Rongeurs by Jegoroff-Freidin N2 -1 pc |  |  |
|  |  | 17. Rongeurs by Jegoroff-Freidin N3 -1 pc |  |  |
|  |  | 18. Rongeurs with semicircular jaws, strong - 1pc |  |  |
|  |  | 19. Rongeurs by Dahlgren for adults - 1 set |  |  |
|  |  | 20. Neurosurgical rongeurs curved over a plane -1 pc |  |  |
|  |  | 21. Neurosurgical rongeurs curved along the edge/rib-1 pc |  |  |
|  |  | 22. Surgical blunt hook, 4-toothed N3 - 2 pcs |  |  |
|  |  | 23. Surgical blunt hook, 4-toothed N4 - 2 pc |  |  |
|  |  | 24. Surgical sharp hook, 3-toothed N1 - 2 pcs |  |  |
|  |  | 25. Surgical sharp hook, 2-toothed - 2 pcs |  |  |
|  |  | 26. Bilateral neurosurgical curette, medium - 1pc |  |  |
|  |  | 27. Bilateral neurosurgical curette, small- 1pc |  |  |
|  |  | 28. Set of oval neurosurgical sharp spoons, rigid- 1set |  |  |
|  |  | 29. Set of round neurosurgical unilateral spoons, sharp- 1set |  |  |
|  |  | 30. Set of oval neurosurgical unilateral spoons, sharp - 1set |  |  |
|  |  | 31. Hook for the insulation of nerve trunks -2 pcs |  |  |
|  |  | 32. Wimble Russian (КННфс-МИЗ-В) or equivalent - 1 set |  |  |
|  |  | 33. Cannulas to puncture the brain ventricles (length 46,66,86,106mm) -2sets |  |  |
|  |  | 34. Removable Clips 7,5 mm - 2000pcs |  |  |
|  |  | 35. Non-removable Clips 7,0 mm - 1000pcs |  |  |
|  |  | 36. Non-removable Clips 4,5 mm - 3000pcs |  |  |
|  |  | 37. Ear Spoon Double-Ended, sharp- 1 pc |  |  |
|  |  | 38. Vertically-curved Scissors with rounded edges , narrow 175mm - 1 pc |  |  |
|  |  | 39 Vertically-curved Scissors, blunt 140mm - 1 pc |  |  |
|  |  | 40. Vascular Scissors, straight 160mm - 1pc |  |  |
|  |  | 41. Horizontally curved vascular Scissors 160mm- 1pc |  |  |
|  |  | 42. Knife for opening the shell of tumor НК210x6,5- 1 pc |  |  |
|  |  | 43. A set of tools for imposing and removing of clips on the brain vessels- 1set |  |  |
|  |  | 44. A set of cutters and drills to wimble- 2 sets |  |  |
|  |  | 45. Twisted saw wire - 50 pcs |  |  |
|  |  | 46. The conductor for wire saws - 5 pcs |  |  |
|  |  | 47. Lifter for soft tissues - 2pcs |  |  |
|  |  | 48. General purpose anatomical tweezers ПА150x2,5 -2pcs |  |  |
|  |  | 49 General purpose surgical l tweezers ПХ 150x2,5 -2pcs |  |  |
|  |  | 50. Toothed gripping tweezers ПХ150x5,5 - 2pcs |  |  |
|  |  | 51. Tweezers for coagulation ПA200x1- pc - 2pcs |  |  |
|  |  | 52. Curved forceps for chiasmal area ПАи200х1,2 -2pcs |  |  |
|  |  | 53. Forceps for dura mater ПХ200х1,3 - 2pcs |  |  |
|  |  | 54 Neurosurgical curved forceps ПАи177х1,5 - 1pc |  |  |
|  |  | 55. Tweezer for holding a brain tumor, fenestrated ПС210х8 - 1 pc |  |  |
|  |  | 56. Tweezer for holding a brain tumor, fenestrated ПС210х17 - 1 pc |  |  |
|  |  | 57. Tweezer for holding a brain tumor, fenestrated ПС210х22 - 1 pc |  |  |
|  |  | 58. Raspatories for spine or chisel with hexagonal handle, flat 40mm (with a rounded spatula) -1 pc |  |  |
|  |  | 59. Handles for wire saws -2 pcs |  |  |
|  |  | 60. Neurosurgical Retractor with sharp jaws -2 pcs |  |  |
|  |  | 61. Raspatories, straight - 2 pcs |  |  |
|  |  | 62. Syringe №1 - 30 cub. cm -1 pc |  |  |
|  |  | 63. Syringe №3 - 90 cub. cm -1 pc |  |  |
|  |  | 64. Ophtalmic scalpel, pointed Co130x20 -2pcs |  |  |
|  |  | 65. Bellied ophtalmic scalpel , medium - Сб150x40 -2pcs |  |  |
|  |  | 66. Neurosurgical Spatula unilateral, with working area width 6mm -2pc |  |  |
|  |  | 67. Neurosurgical Spatula unilateral, with working area width 10mm -2pc |  |  |
|  |  | 68. Neurosurgical Spatula unilateral, with working area width 20mm -2pc |  |  |
|  |  | 69. Neurosurgical Spatula unilateral, with working area width 26mm -2pc |  |  |
|  |  | 70. Neurosurgical Spatula bilateral, with working area width 8mm -2pc |  |  |
|  |  | 71. Neurosurgical Spatula bilateral, with working area width 15mm -2pc |  |  |
|  |  | 72. Neurosurgical Spatula bilateral, with working area width 20mm -2pc |  |  |
|  |  | 73. Double-sided elevator for separation soft tissues from nasal osteochondral septum wall - 1pc |  |  |
|  |  | 74. The spoon-shaped forceps to remove a brain tumor - 1 pc. |  |  |
|  |  | 75. Fenestrated forceps for grasping the brain tumor - 1 pc. |  |  |
|  |  | 76. Osteal forceps, nippers hinged,with double transfer and straight jaws, vurved in a plane -1 pc. |  |  |
|  |  | 77. Protective forceps for drilling the skull-1pc |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 29 | Վիրահատական գործիքների հավաքածու ինտրամեդուլյար մեխ դնելու համար // Set of instruments for installing intramedullary nails for femoral bone with case (universal) | NOTE: the following specifications identify the desired configuration of the required Set of instruments for installing intramedullary nails for femoral bone with case (universal) and of its components. Bidders are allowed to offer functionally equivalent solutions, including all components described below. | 1 |  |
|  |  | 1.The bracket for fixing intramedullary nails, consisting of: - conductor assembly handle - 1 pc. (Handle - 1 pc, the pilot - 1 pc., screw M6 -. 1 pc) - pilot proximal - 1pc - constricting bolt - 1pc - cap-anvil - strip of conductor for cervical and femoral nails (for 5 apert.) - 1 pc - cervical strip of cundoctor for the reconstruction of nails (for 4 holes) - 1pc - cervical strip of cundoctor for the proximal nails (for 2 holes). - 1 pc - diaphyseal panel of conductor for proximal nails (for 2 holes). - 1 pc - rod for adjustment the conductor, diameter 4,9mm - 1pc - rod for adjustment the conductor, diameter 6,0 mm - 1pc - rod for adjustment the conductor, diameter 6,5 mm - 1pc - rod for adjustment the conductor, diameter 11 mm - 1pc |  |  |
|  |  | 2.Set of trocar with guides, including: - troacar, diameter 8mm - 1pc - guide, diameter 12 - 8 mm - 1pc - guide, diameter 14 - 12 mm - 1pc - guide, diameter 11 - 8 mm - 1pc - guide, diameter 8- 5 mm - 1pc - guide, diameter 8 - 4 mm - 1pc - guide, diameter 8 - 2.5 mm - 1pc |  |  |
|  |  | 3. Screwdriver under inner hex S=3,5; L = 350mm with flat metallic handle - 1pc |  |  |
|  |  | 4. Extractor percussion - 1 pc |  |  |
|  |  | 5. Chisel cannulated, composed of: - chisel - 1pc - protector - 1 pc |  |  |
|  |  | 6. Hammer slit-shaped, including : - screwdriver - 1pc - Guide - 1 pc |  |  |
|  |  | 7. Screwdriver for proximal screw L=342mm, including : - screwdriver - 1pc - constricting bolt - 1pc |  |  |
|  |  | 8. Channel depth measurer L=100mm - 1pc |  |  |
|  |  | 9. Case |  |  |
|  |  | 10. Combination wrench 11 x 11 mm - 1 pc |  |  |
|  |  | 11. Drill diam.4,0 L = 310 mm, stainless steel. - 1pc |  |  |
|  |  | 12. Drill diam.5,0 L = 310 mm, stainless steel. - 1pc |  |  |
|  |  | 13. Drill diam.6,5 L = 310 mm, stainless steel. - 1pc |  |  |
|  |  | 14. Drill diam.11 L = 310 mm, stainless steel. - 1pc |  |  |
|  |  | Additional tools: |  |  |
|  |  | Navigator |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 30 | ԼՈՌ էնդոսկոպիկ վիրահատությունների գործիքների հավաքածու // Surgical ENT set | NOTE: the following specifications identify the desired configuration of the required Surgical ENT set and of its components. Bidders are allowed to offer functionally equivalent solutions, including all components described below. | 1 |  |
|  |  | 1 . Adenotome №1 - 1pc |  |  |
|  |  | 2 . Adenotome №3 - 1pc |  |  |
|  |  | 3. Adenotome №5 - 1pc |  |  |
|  |  | 4. Ear Specula nickel-plated No. 1 - 1pc |  |  |
|  |  | 5. Ear Specula nickel-plated No. 2 - 1pc |  |  |
|  |  | 6. Ear Specula nickel-plated No. 3 - 1pc |  |  |
|  |  | 7. Ear Specula nickel-plated No. 4 - 1pc |  |  |
|  |  | 8. Chisel with fluted handle, curved 4mm - 1pc |  |  |
|  |  | 9. Chisel with fluted handle, curved 6mm - 1pc |  |  |
|  |  | 10. Chisel with fluted handle, flat 4mm - 1pc |  |  |
|  |  | 11. Chisel with fluted handle, flat 6mm - 1pc |  |  |
|  |  | 12. Hemostatic forceps, grooved straight N1 - 2 pcs |  |  |
|  |  | 13. Hemostatic forceps, grooved straight N2 - 2 pcs |  |  |
|  |  | 14. Laryngeal speculum diameter 15 mm - 1pc |  |  |
|  |  | 15. Laryngeal speculum diameter 22 mm - 2pcs |  |  |
|  |  | 16. Nasopharyngeal Speculum, diameter 8mm - 1pc |  |  |
|  |  | 17. Nasal Speculum, jaw length 40mm - 1pc |  |  |
|  |  | 18. Nasal probe by Voyachek, bellied - 3pcs |  |  |
|  |  | 19. Ear probe by Voyachek, bellied -3pcs |  |  |
|  |  | 20. Nasal probe with wrapping |  |  |
|  |  | 21. Needle by Kulikovsky to puncture the maxillary sinus - 1pc |  |  |
|  |  | 22. Aural catheter, metallic КУ-2 - 1pc |  |  |
|  |  | 23. Aural catheter, metallic КУ-3 - 1pc |  |  |
|  |  | 24. Conchotome with slit-shaped hole N1 or N2 - 1pc |  |  |
|  |  | 25. Tracheotomy hook, sharp- 1pc |  |  |
|  |  | 26. Hook for pulling the wings of the nose - 2pcs |  |  |
|  |  | 27. Hook for the extraction of foreign bodies from the nose 115 mm - 1pc |  |  |
|  |  | 28. Hook for the extraction of foreign bodies from the ear 125 mm - 1pc |  |  |
|  |  | 29. Surgical hook sharp, 3-toothed N1 - 2pcs |  |  |
|  |  | 30. Ear Spoon sharp , big, tough - 1pc |  |  |
|  |  | 31. Surgical Hammer wooden - 1pc |  |  |
|  |  | 32. Scissors for dissecting nasal turbinates, horizontal curved 160 mm - 1 pc |  |  |
|  |  | 33. Surgical scissors straight blunt, 140 mm - 1pc |  |  |
|  |  | 34. Loop polypouses for ear (set.) - 1pcs |  |  |
|  |  | 35. Loop polypouses nasal (set.)- 1pcs |  |  |
|  |  | 36. Anatomical tweezers ПА 150х2,5 - 2 pcs |  |  |
|  |  | 37. Crampon forceps for dental works and splinting of jaws - 1 pc |  |  |
|  |  | 38. Wire 2-0,2 - 6 hanks |  |  |
|  |  | 39. Wire 2-0,3 - 9 hanks |  |  |
|  |  | 40. Tracheotomy dilator - 1 pc |  |  |
|  |  | 41. Frontal reflector РМО - 1pc |  |  |
|  |  | 42. Scalpel Сб 150х40 - 1pc |  |  |
|  |  | 43. Scalpel Со 150х40 - 1pc |  |  |
|  |  | 44. Knife (needle) НК 175х4 bayonet type paracentetic - 1pc |  |  |
|  |  | 45. Tonsillectomy (set ) - 1pcs |  |  |
|  |  | 46. Tracheotomy tube, plastic N 1 - 2 pcs |  |  |
|  |  | 47. Tracheotomy tube, plastic N 3 - 2 pcs |  |  |
|  |  | 48. Tracheotomy tube, plastic N 3 - 2 pcs |  |  |
|  |  | 49. Tracheotomy tube, plastic N 5 - 2 pcs |  |  |
|  |  | 50. Spatula for tangue, double-sided, straight - 5 pcs |  |  |
|  |  | 51. Syringe for rinsing the cavities with double scale - 1pc |  |  |
|  |  | 52. Syringe for endolaryngeal infusions and rinsing of the tonsils capacitance 2mm - 2 set |  |  |
|  |  | 53. Laryngeal forceps to retrieve foreign bodies, curved 200 mm - 1 pc |  |  |
|  |  | 54. Forceps for operation on nasal septum - 1 pc |  |  |
|  |  | 55. Nasal Dressing Forceps - 1 pc |  |  |
|  |  | 56. Ear Dressing Forceps - 1 pc |  |  |
|  |  | 57. Mastoid Rongeur, straight - 1pc |  |  |
|  |  | 58. Mastoid Rongeur, curved - 1pc |  |  |
|  |  | 59. Bone Forceps-nippers with round jaws, straight - 1 pc |  |  |
|  |  | 60. Elevator for tonsils - 1pc |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 31 | Աուդիոմետեր //Audiometer | Should have not less than a diagnostic 2 channel audiometer | 1 |  |
|  |  | Should have capable of assessing full range of frequencies including high frequency audiometry from 125 to 12 000 Hz |  |  |
|  |  | Should have intensity range of from -10 to 120 for air conduction, bone conduction from -10 to 75, sound field from -10 to 90 dB |  |  |
|  |  | Must have air conduction and bone conduction |  |  |
|  |  | Should have capable of tone testing |  |  |
|  |  | Should have headphones and inserts connections separate, allowing switching to either without having to change plug ins. |  |  |
|  |  | Should be PC compatible, possibility of printer connection |  |  |
|  |  | Should have automatic data recovery in the event of power outage |  |  |
|  |  | Should have capable of test storage |  |  |
|  |  | Should have capable of generating report |  |  |
|  |  | Should have capable of completing an electronic record |  |  |
|  |  | Calibration: Equipment with calibrations accordind to ISO and ANSI |  |  |
|  |  | Working temperature: from 10 to 40oC |  |  |
|  |  | Аccessories |  |  |
|  |  | The set includes all the necessary additional accessories to complete functioning of the equipment (example computer and ect) |  |  |
|  |  | Power Supply |  |  |
|  |  | Power input: 220 to 240V/ 50 Hz |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 32 | Նորածինների սիպապ / CPAP machine | CPAP machine | 1 |  |
|  |  | CPAP generator with pressure range from 3 to 10 cm of water. Should haveCapable of giving nasal/nasopharyngeal CPAP. |  |  |
|  |  | Integrated pneumatic Air and oxygen blender calibrated with flow from 0 to 15 lit/min. |  |  |
|  |  | Safety mechanism for relief of excessive pressure through pressure relief valve/regulator. |  |  |
|  |  | Soft nasal prongs |  |  |
|  |  | Should have alarms for the device: Low/high Temperature, Flow increase/decrease alarm, O2 pressure low alarm, Air pressure low alarm |  |  |
|  |  | Flow meters: 02 with each piece. |  |  |
|  |  | Power pack (UPS with battery back of minimum 1 hour). |  |  |
|  |  | Power Supply |  |  |
|  |  | Power input: 220 to 240V/ 50 Hz |  |  |
|  |  | Accessories |  |  |
|  |  | Heated wire servo-controlled humidifier - 1pc |  |  |
|  |  | 5 ml test lung - 1pc |  |  |
|  |  | Disposable patient circuits - 30pcs |  |  |
|  |  | Disposable nasal prongs - 30pcs (10 each of different neonatal sizes) |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 33 | Նորածինների արտածծիչ // Newborn suction pump | Portable suction machine for minor surgery, professional use, with powerful piston motor. | 1 |  |
|  |  | Ideal for home post-surgery treatment |  |  |
|  |  | Anti-shock case, light, easy to carry and keep clean. |  |  |
|  |  | Supplied with 1 lt. jar serializable, with over flow valve. |  |  |
|  |  | Complete with suction regulator, vacuum gauge. |  |  |
|  |  | Power consumption: 160W, deviation not more than 10% |  |  |
|  |  | Suction power: not less than 35 l/min |  |  |
|  |  | Max vacuum: not less than 0.8 bar |  |  |
|  |  | Jar: 1x1L |  |  |
|  |  | Accessories |  |  |
|  |  | Set of filter disposable |  |  |
|  |  | Set of disposable tubes |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 34 | էնդովենոզ լազեր // Endovenous laser | Endovenouslazer treatment of greater and lesser varicose saphenous veins | 1 |  |
|  |  | Should have in bilt power meter |  |  |
|  |  | Minimum risk of side effects |  |  |
|  |  | Improved aesthrtic results |  |  |
|  |  | Nopost op downtime |  |  |
|  |  | Fast outpactient procedure |  |  |
|  |  | Need local anaesthrsia only |  |  |
|  |  | Output power range: 980nm, deviation not more than 5% |  |  |
|  |  | Optional should have 1470nm , deviation not more than 5% |  |  |
|  |  | Aming beam: through fibre 635nm , deviation not more than 5% |  |  |
|  |  | Mode of operation: continuous wave or pulsed |  |  |
|  |  | Frequency range: from 0.3 to 100Hz , deviation not more than 5% |  |  |
|  |  | Pulse width: from 10 to 300ms , deviation not more than 5% |  |  |
|  |  | Lazer type: continuous wave diode lazer not less than Class 4 (IEC825) |  |  |
|  |  | Accessories |  |  |
|  |  | The set includes all the necessary additional accessories to complete functioning of the equipment |  |  |
|  |  | Power Supply |  |  |
|  |  | Power input: 220 to 240V/ 50 Hz |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 35 | Հեմոդիալիզի սարք և պարագաներ //Haemodialysis Machine and accessories | Operational Requirements | 15 |  |
|  |  | Machine should have facility for Acetate, Bicarbonate, Sequential dialysis (Isolated UF) |  |  |
|  |  | Upgradable to future software development and linkable with Patient Data Management System |  |  |
|  |  | The blood pump should be able to run from 15 to 650 ml per min and adaptable to standard A -v blood lines and should run even in the absence of water or dialysis flow |  |  |
|  |  | Technical Specifications |  |  |
|  |  | Should have facility for conventional and high flux dialysis. |  |  |
|  |  | Machine should have bacterial filter (Pyrogen filter) before water going to dialyser. |  |  |
|  |  | Battery back-up for temporary maintaining of extracorporeal blood circulation. |  |  |
|  |  | Should have Na and UF profiling |  |  |
|  |  | Dialysate temperatures selectable between 35 degrees C to 39 degrees C. |  |  |
|  |  | Variable conductivity setting between 12.8 to 15.7. |  |  |
|  |  | Should have variable dialysate flow 300 - 800 ml/mt and should have increasing facility to step up by 100 ml., or AutoFlow. |  |  |
|  |  | Should have facility to show trends. |  |  |
|  |  | Heparin pump with syringes sizes up to 50 ml with pump flow rate from 0.1-10 ml per hr |  |  |
|  |  | (0.1 ml increments). |  |  |
|  |  | Stroke pressure operated short term single needle dialysis |  |  |
|  |  | Ultrafiltration 0.1 to 2.5 litres per hr. The in and out fluid circuit must be separated so that there is no chance of contamination in the event of membrane rupture. |  |  |
|  |  | Treatment parameter should be displayed by graph and digitally both. |  |  |
|  |  | Should have integrated heat and chemical disinfection facility with both short and long disinfection programme with day night week schedule. |  |  |
|  |  | Should have accurate feedback control conductivity mixing technique. |  |  |
|  |  | Should have drain facility. |  |  |
|  |  | Should have accurate UF control by flow by volume control measurement technique. |  |  |
|  |  | Extra facilities like online clearance monitring |  |  |
|  |  | Presetable important data so that machine can be used anytime without feeding data every time. |  |  |
|  |  | Should have automatic self test facility |  |  |
|  |  | Should have auto on/ off facility |  |  |
|  |  | Should have touch button screen and large colour TFT screen. |  |  |
|  |  | Automatic diagnosis of malfunctioning with on line ability to show the faults with trouble shooting (technical service mode). |  |  |
|  |  | Machine can be connected to computer to feed data and trouble shoot whenever problem arises. |  |  |
|  |  | Audio visual alarms on limit violation of conductivity, blood leak, air leak, transmembrane pressure alarms, Dialysis temperature alarm, end of disinfection alarm, bypass alarm and blood pump stop alarm. |  |  |
|  |  | On line creatinine clearance monitoring -build in device for measurement and monitoring of effective urea clearance and dialysis dose (KT/V) |  |  |
|  |  | System configuration Accessories, spare and consumables. |  |  |
|  |  | All consumables required for installation and standardization of system to be given. |  |  |
|  |  | The set includes all the necessary additional accessories to complete functioning of the equipment |  |  |
|  |  | Power Supply |  |  |
|  |  | Power input: 220 to 240V/ 50 Hz |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
|  |  | Հեմոդիալիզի բազաթոռ // Haemodialysis bed |  |  |
|  |  | Should be ergonomically designed and comfortable to the patient. |  |  |
|  |  | Should allow the patient to try in full sitting and lying position |  |  |
|  |  | Should have electronically controlled adjustment for back section, leg section and height. |  |  |
|  |  | Should have a patient hand set with controls for all positions. |  |  |
|  |  | Arm set should fold to allow side entry of the patient |  |  |
|  |  | Seat cushion should be removable, made of proper density foam and should have smooth surface for easy hygiene and cleaning |  |  |
|  |  | Frame should be made up of corrosion free galvanized steel with powder coating and should have four 150mm dia swiveling castor wheels of which the front two should be lockable. |  |  |
|  |  | Should be able to withstand a maximum load of 150kg |  |  |
|  |  | Should have facility for weight measurement |  |  |
|  |  | Should have detachable drip stand and a tray table |  |  |
|  |  | Power input: 220 to 240V/ 50 Hz |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |
| 36 | Ձրի մաքրման կայան հեմոդիալիզի սարքավորումների համար //Water treatment plant for hemodialysis equipment | Water Treatment Plant (WTP) with capacity 1000 Litre/Hour Post RO water | 1 |  |
|  |  | It should have capacity to produce 1000 Litre/Hour post RO water. |  |  |
|  |  | Should be of Microprocessor based RO system |  |  |
|  |  | It should have Raw water tank, Multi grade Filter, Iron removal, Charcoal Filter, softener, Reverse Osmosis, etc. and should have internal plumbing for post RO water supply to each machine by PEX pipeline with character to sustain heat up to 93 degree centigrade for heat disinfection. (The pre-Treatment and RO System should be fully automatic) |  |  |
|  |  | It should have Rines and Flush facility for the membrances in R.O. Unit. |  |  |
|  |  | It should have Conductivity meter, Pressure Gauge, Flow indicator. |  |  |
|  |  | Should have BIS standard/for assembled component the quality standard should be as per approval of appropriate standard authority related to used component |  |  |
|  |  | It should have Stainless Steel Connectors for Water Outlet at Dialysis machine connection points |  |  |
|  |  | The Chemical, contaminants, endotoxin and microbial tests of post R.O. water should pass AAMI standards. |  |  |
|  |  | The Chemical, contaminants, endotoxin and microbial tests of post R.O. water should pass AAMI standards. The pipeline for RO water supply along with the connectors and the drainage pipe materials will be supplied by the RO vender to the person responsible for the Turn key project of the Dialysis unit. The maintenance of the pipeline will be done by RO supplier |  |  |
|  |  | Should have display for supply and temperature of Permeate |  |  |
|  |  | System configuration Accessories, spare and consumables. |  |  |
|  |  | The set includes all the necessary additional devices and accessories to complete functioning of the equipment (example: additional water pump, water tank and ect. ) |  |  |
|  |  | Power Supply |  |  |
|  |  | Power input: 220 to 240V/ 50 Hz |  |  |
|  |  | Note: REQUIRED TECHNICAL SERVICES. Provide services of a qualified manufacturer's representative to check complete installation for conformance to manufacturer's recommendations, put system into service, make all adjustments required for full conformance to design and specified requirements, and perform all demonstrations and tests. |  |  |
|  |  | Note: STARTUP AND TESTING. Tests shall be run in presence of Contracting Officer's Technical Representative (COTR) or Resident Engineer (RE). |  |  |
|  |  | Note: The system performance should be designed for operation of 20 dialysis machine. |  |  |
|  |  | Note: All changes, during the installation, related to ensuring operability of the equipment must be provided by the own means of the supplier. |  |  |
|  |  | Standards (Documentary evidence (copy of certificates): |  |  |
|  |  | ISO13485 or equivalent |  |  |
|  |  | CE Mark (Directive 93/42/EEC) or equivalent |  |  |

Technical specifications

Lot 4 - Medical Furniture

| N | NAME | TECHNICAL SPECIFICATION REQUIRED BY THE PURCHASER | Q-ty |  |  |
| --- | --- | --- | --- | --- | --- |
| 1 | Տեղեկատույի սեղան // Reception desk | Plastic laminated consisted of 3 wings. | 1 |  |  |
|  |  | Size of the main desk: height/length/width – 1100mm/5000mm/450mm, deviation not more than 10% |  |  |  |
|  |  | On the right side of each work desk 4 drawers with Nylon ball bearing Sliders, one with locker, height -150mm, width - 400mm |  |  |  |
|  |  | Downward-projecting Lamp fixed on main desk with accessible switch. The Purchaser can ask the changes of the dimensions during the contract implementation, in case such changes will not create additional costs for the Supplier. |  |  |  |
| 2 | Տեղեկատույի սեղան ամբուլատոր բաժանմունքի համար // Reception desk for outpatient department | Plastic laminated consisted of 2wings. | 1 |  |  |
|  |  | Size of main desk: height/length/width – 1100mm/3800mm/450mm. deviation not more than 10% |  |  |  |
|  |  | On the right side of each work desk 4 drawers with Nylon ball bearing Sliders, one with locker, height -150mm, width - 400mm |  |  |  |
|  |  | Downward-projecting Lamp fixed on main desk with accessible switch. The Purchaser can ask the changes of the dimensions during the contract implementation, in case such changes will not create additional costs for the Supplier. |  |  |  |
| 3 | Տեղեկատուի սեղան շտապ օգնության բաժանմունքի համար // Reception desk for emergency | Plastic laminated consisted of 2wings. | 1 |  |  |
|  |  | Size of main desk: height/length/width – 1100mm/5300mm/450mm. deviation not more than 10%.Plastic laminated consisted of 2 parts: main desk and rectangular attachments. |  |  |  |
|  |  | On the left side of work desk should be one drawer with door for filing, height 450mm and width 400mm for filing, and one shelf with height 150mm and width 400mm. |  |  |  |
|  |  | Downward-projecting Lamp fixed on main desk with accessible switch. The Purchaser can ask the changes of the dimensions during the contract implementation, in case such changes will not create additional costs for the Supplier. |  |  |  |
| 4 | Պլաստիկ աթոռ // Plastic Chairs | Fixed plastic chair 400x450mm, deviation not more than 10%, with backrest. | 520 |  |  |
|  |  | Tubular steel frame. Height of seat: from 400 to 500mm |  |  |  |
| 5 | Գրասեղան // Writing Desk | Laminated Top. | 80 |  |  |
|  |  | Single Pedestal |  |  |  |
|  |  | Dimensions: |  |  |  |
|  |  | Height: 700mm, deviation not more than 10% |  |  |  |
|  |  | Length: 1400mm, deviation not more than 10% |  |  |  |
|  |  | Width: 700mm, deviation not more than 10% |  |  |  |
|  |  | 3 drawers and one shelf totally with 600mm height, length of each 400mm |  |  |  |
|  |  | 3 drawers with nylon ball bearing sliders and one with lock. |  |  |  |
| 6 | Պահարան // Cabinet | Plastic Laminated frame Cabinet , 2 Doors with Locks | 45 |  |  |
|  |  | Dimensions: |  |  |  |
|  |  | Height: from 1800 to 1900mm |  |  |  |
|  |  | Length: from 800 to 900mm |  |  |  |
|  |  | Width: from 450 to 550mm |  |  |  |
|  |  | 5 Shelves |  |  |  |
| 7 | Բազկաթոռ // Chair | Washable. Tubular steel frame. Seat and backrest from leather-cloth upholstered. Revolving typing chair, 5 casters base, with armrest, adjustable height with gas pump. | 90 |  |  |
|  |  | Dimensions: |  |  |  |
|  |  | Height: 600mm, deviation not more than 10% |  |  |  |
|  |  | Length: 410mm, deviation not more than 10% |  |  |  |
|  |  | Width: from 400 to 500 mm |  |  |  |
| 8 | Զգեստապահարան (պալատային կահույք) // Wardrobe | Length \ height \ depth (width) | 78 |  |  |
|  |  | 1500...1600mm x 1960...1980mm x 500...520 mm |  |  |  |
|  |  | Fully constructed of plastic laminated 18 mm in thicknees, with two door cabinets on the right side, 20 sm fare from upper and lower levels. |  |  |  |
|  |  | Indoor a hanger prepared to be nickel plated tube. Wardrobe on the left side of the upper and lower parts of the two and one single-door cabinet 60 cm in height. |  |  |  |
|  |  | All doors made ​​of glossy profile and woodchip board. |  |  |  |
|  |  | All Shelves made ​​of laminated woodchip board. |  |  |  |
|  |  | Backside of the wardrobe, covered with sinks fibreboard (ДВП) -of 2.5-3 mm thicknees, with cut for the refrigerator. |  |  |  |
|  |  | Cabinet legs are nickel-plated, 120 mm in height. |  |  |  |
| 9 | Պալատային մահճակալ // Bed With Backrest | Two rocker manual hospital bed | 146 |  |  |
|  |  | Function (not less than present functions or analog): lifting, leg bending, whole bed HI-LO adjustment |  |  |  |
|  |  | Bed plate size: 2140x960x500mm, deviation not more than 10% |  |  |  |
|  |  | Backreast angle adjustment: from 0o to 85o, deviation 5o |  |  |  |
|  |  | Leg plate angle adjustment: from 0o to 45o, deviation 5o |  |  |  |
|  |  | Free HI-LO controls offolded-type guard rails, safe and reliable, easy opeartions |  |  |  |
|  |  | High quality carbon-steel is applied for the main body of the bedstead, all surface antirust treatment, enamel paint of antibacterial powder |  |  |  |
|  |  | One off impact molding of cold-rolled sheet for bed plate, spraying paint of static power, impact-resistant and antirust |  |  |  |
|  |  | The bed is mounted on four 125 mm castors, two fixed at head end and two swiveling with independent brakes at foot end. |  |  |  |
|  |  | Bumpers at all 4 corners. |  |  |  |
|  |  | Include accessories: |  |  |  |
|  |  | Mattress with high-density stretchy sponge, cover withwaterproof cloth only. Overall size L/W/H: 2140x960x500mm, deviation not more than 10% |  |  |  |
| 10 | Ֆունկցիոնալ մահճակալ // Rehabilitation Bed (electric bed) | Rehabilitation Bed (electric bed) | 35 |  |  |
|  |  | Electric bed: not less than 2 function |  |  |  |
|  |  | Function (not less than present functions or analog): back lifting, leg bending, HI-LO adjustment, bed head/end titling, head board, foot board: ABS integrated forming |  |  |  |
|  |  | Bed plate size: 1930x850x450/770mm, deviation not more than 10% |  |  |  |
|  |  | High quality carbon-steel is applied for the main body of the bedstead, all surface antirust treatment, enamel paint of antibacterial powder |  |  |  |
|  |  | One off impact molding of cold-rolled sheet for bed plate, spraying paint of static power, impact-resistant and antirust |  |  |  |
|  |  | Full enclosed permanent magnet DC motor, strong thrust, low noise in operation, waterproof and dustproof UPS, making nursing more reliable and professional |  |  |  |
|  |  | 4 central controlling casters, great bearing capacity, flexible rotation, stable and reliable |  |  |  |
|  |  | The bed is mounted on four 125 mm castors, two fixed at head end and two swiveling with independent brakes at foot end. |  |  |  |
|  |  | 4 Section Mattress Platform fitted with epoxy- coated steel weldmesh backrest and knee brake adjustable by electro-hydraulic action. |  |  |  |
|  |  | Bumpers at all 4 Corners |  |  |  |
|  |  | Include accessories: Mattress with high-density stretchy sponge, cover withwaterproof cloth only. Overall size L/W/H: 1930x850x450mm, deviation not more than 10%, IV Rod,, Side Folding-down rails (Pair). |  |  |  |
| 11 | Պալատային կիսապահարան // Bedside cabinet | Material: ABS engineering plastic. | 193 |  |  |
|  |  | Towel racks and hooks on both sides to hold light things. |  |  |  |
|  |  | Extendable table board, adjustable laminates in the cabinet |  |  |  |
| 12 | Մահճակալային սեղան // Movable table | Material: Metal | 193 |  |  |
|  |  | Laminated board for table top, leg supports with high quality carbon-steel surface spraying painted, electroplated internal sliding tube |  |  |  |
|  |  | Controllable air spring height adjustment. |  |  |  |
|  |  | Gray rubber controllable casters. |  |  |  |
| 13 | Կիսապահարան // Bedside Locker | Bedside Locker or Cabinet for Patient Bedroom accommodation | 46 |  |  |
|  |  | Shall be fully constructed from laminated Plywood, |  |  |  |
|  |  | 450/550mm, deviation not more than 10% , |  |  |  |
|  |  | Height: 700mm, deviation not ore than 10% |  |  |  |
|  |  | Top gallery or embossed plastic recess for objects safety |  |  |  |
|  |  | Shall be delivered as kit to be assembled |  |  |  |
|  |  | Shall have 2 Castors (Orientale) and 2 Feet with rubber Studs |  |  |  |
|  |  | Upper Drawer: 450/200mm , deviation not more than 10% ,with Nylon ball bearing Sliders. |  |  |  |
|  |  | Lower drawer: 450/400mm , deviation not more than 10% |  |  |  |
| 14 | Մանկական մահճակալ // Baby bed | Baby bed | 6 |  |  |
|  |  | Durable frame: Made of cold-rolled steel plate and tube, coated by electrostatic spray after twice phosphatization. |  |  |  |
|  |  | Steel bedboard |  |  |  |
|  |  | Foldable al-alloy handrails |  |  |  |
|  |  | With silent wheels with cross brakes |  |  |  |
|  |  | Fully tested sturdy contruction capable of taking a maximum user weight of up to 200kg |  |  |  |
|  |  | Matress: availability |  |  |  |
|  |  | Technical Parameters: |  |  |  |
|  |  | Overall size (LxWxH): 1210x690x600mm, deviation not more than 10% |  |  |  |
| 15 | Նորածնի մահճակալ // Baby crib | Trolley manufactured from rust-proofed steel tube mounted on four casters ∅60mm. | 5 |  |  |
|  |  | Crib manufactured from Plexiglas, with aeration holes and with position TR/RTR. |  |  |  |
|  |  | Length: 825mm, deviation not more than 10% |  |  |  |
| 16 | Կախիչ // Hanger | Coat rack without doors, laminated plank. | 75 |  |  |
|  |  | Metallic hooks 3 or 4 unit |  |  |  |
|  |  | Dimensions: |  |  |  |
|  |  | Height: 600mm, deviation not more than 10% |  |  |  |
|  |  | Length: 600mm, deviation not more than 10% |  |  |  |
| 17 | Պարկով թափոնների սայլակ // Soled linen trolley | Chrome plated frame. Mounted on Ø 75 mm casters. Removable bag. Knock down construction. 1 Bag. 750mmx550mmx550mm, deviation not more than 10% | 6 |  |  |
| 18 | Թափոնների սայլակ // Bowl`s stand | Constructed from chrome plated steel tube. Variable height. | 6 |  |  |
|  |  | Stainless Steel bowl. Single |  |  |  |
|  |  | BOWL - Stainless steel 18/10, Ø 32 cm. |  |  |  |
|  |  | Capacity: not less than 4L |  |  |  |
|  |  | Stainless steel base. 5 castor base ø 50 mm. |  |  |  |
| 19 | Լաբորատոր սարքավորումների սեղան //Table for laboratory equipment | Workbench for storage of laboratory equipment and procedures such as Tests / Examinations without use of Water ( Dry Workbench ) | 7 |  |  |
|  |  | Bench Top shall be made from MDF, thickness not less than 36mm, laminated on both sides (Counterbalanced). Laminate thickness not less 18 mm. Bench Top shall match “Humid Bench Tops “in the Laboratory and for that purpose Bench Top shall be made from one piece. |  |  |  |
|  |  | Dimensions not less 1500 x750 x 700mm. |  |  |  |
|  |  | Shelving shall be constructed with Laminated Plywood on 2 Levels: 1st Level at 400mm and 2nd Level at 700mm from Bench Top. Widths shall be respectively 150 and 300 mm (Top Shelf). Shelves shall be supported by lateral and walled-panels from identical materials |  |  |  |
|  |  | The Upper Shelf shall be equipped with downward-projecting Lamp with at least 600 mm length, with accessible switch. |  |  |  |
|  |  | The Workbench shall be equipped with 4 Grounded Electrical Socket Outlets, 10A. situated on the Wall Panel at 950mm FFL ( From Finished Floor ) |  |  |  |
|  |  | Lower Cupboard with 2x500mm width, height: 800mm±50mm Doors (Large opening Capability), Laminated construction, 2 Laminated Shelves. Handles and Locks. May be optionally made from Steel and independent from the Bench |  |  |  |
|  |  | Set of 3 Drawers |  |  |  |
|  |  | 400 mm wide Laminated material , 2 small with 150mm±10%, height and 1 Large with 300mm±10%, height with silent ball. Bearings Sliders with end safety , Handles and Lock only for the large drawer. |  |  |  |
| 20 | Հետազոտման թախտ // Medical examination couch | Couch for Patient Examination, shall be 2 Section Type (Backrest ) | 40 |  |  |
|  |  | Adjustable Backrest thru manual Ratchets made of chrome-plated steel, Action onboth sides rigid top , Mattress with anti-fire PVC & sponge |  |  |  |
|  |  | Upholstered simile-Leather, Washable CoverChrome-Plated heavy duty tubular Frame with provision for VisesLongitudinal strengthening Bar. Floor protection by Rubber Studs. |  |  |  |
|  |  | Dimensions: |  |  |  |
|  |  | Length: 1850 mm, deviation not more than 10% |  |  |  |
|  |  | Width: 620mm, deviation not more than 10% |  |  |  |
|  |  | Height: 750 mm, deviation not more than 10% |  |  |  |
| 21 | Բժշկական պահարան // Cabinet for medicine | Intended for storage of DrugsShall be constructed from laminated MDF or Sheet Steel protected with bakedEpoxy Powder coating supported by legs with adjustable height Studs. Plinth maybe optionally proposed | 65 |  |  |
|  |  | Shall have 2 Leaves Door with Handles and Locks |  |  |  |
|  |  | Shall have 4 height shelves |  |  |  |
|  |  | Shall have a toxics storage Compartment, Red Color , with Locks |  |  |  |
|  |  | Dimensions: |  |  |  |
|  |  | Depth: 400 mm, deviation not more than 10% |  |  |  |
|  |  | Length: 600 mm, deviation not more than 10% |  |  |  |
|  |  | Height: 1600 mm, deviation not more than 10% |  |  |  |
| 22 | Հեռուստացույց պատի կախիչով (Գունավոր հեռուստացույց-43" ՚՚ամրակներով) // Color LED TV - 43"with wall holder | Color LED TV Set. Russian/English OSD | 3 |  |  |
|  |  | Size: not less 43" |  |  |  |
|  |  | Automatic/semi-automatic tuning |  |  |  |
|  |  | On-screen channels indication |  |  |  |
|  |  | Direct program choice |  |  |  |
|  |  | TV set Remote control |  |  |  |
|  |  | 2 batteries |  |  |  |
|  |  | Internal block for digital TV |  |  |  |
|  |  | Power: 220V/50 Hz. deviation not more than 10% |  |  |  |
|  |  | Metal frame holder affix on wall |  |  |  |
| 23 | Հեռուստացույց տակդիրով (Գունավոր հեռուստացույց 32" ՚՚ամրակներով) // Color LED TV - 32"with holder | Color LED TV Set. Russian/English OSD | 78 |  |  |
|  |  | Size: not less 32" |  |  |  |
|  |  | Automatic/semi-automatic tuning |  |  |  |
|  |  | On-screen channels indication |  |  |  |
|  |  | Direct program choice |  |  |  |
|  |  | TV set Remote control |  |  |  |
|  |  | 2 batteries |  |  |  |
|  |  | Internal block for digital TV |  |  |  |
|  |  | Power: 220V/50 Hz. deviation not more than 10% |  |  |  |
|  |  | Metal frame holder affix on wall |  |  |  |
| 24 | Սառնարան պալատային // Refrigerator | Capacity- not less than 100 liters | 83 |  |  |
|  |  | Number of shelves: not less than 2 |  |  |  |
|  |  | Power: 220V/50 Hz. deviation not more than 10% |  |  |  |
|  |  | Euro standard plug |  |  |  |
|  |  | Washable surface. |  |  |  |
| 25 | Դարակաշար // Cabinet with Shelves | Laminated cabinet with: | 75 |  |  |
|  |  | Dimensions: |  |  |  |
|  |  | Length: 1000mm, deviation not more than 10% |  |  |  |
|  |  | Height: 1800 mm, deviation not more than 10% |  |  |  |
|  |  | Width: 450mm, deviation not more than 10% |  |  |  |
|  |  | Should have 3 adjustable shelves: |  |  |  |
|  |  | Length: 1000mm, deviation not more than 10% |  |  |  |
|  |  | Width: 450mm, deviation not more than 10% |  |  |  |
|  |  | Load bearing capacity not less 50kg. |  |  |  |
| 26 | Սեղան պալատային // Table | Laminated top, chrome plated legs. | 100 |  |  |
|  |  | Dimensions: |  |  |  |
|  |  | Height: from 700 to 750 mm. |  |  |  |
|  |  | Length: 500 mm, deviation not more than 10% |  |  |  |
|  |  | Width: 800 mm, deviation not more than 10% |  |  |  |
| 27 | Քույրական պոստի սեղան#1 // Nurse station#1 | Registration desk of laminate 18mm thick | 8 |  |  |
|  |  | Sizes: length/width/height- 1400…1600mm/700…750mm/750mm. |  |  |  |
|  |  | Station should have 2 drawers, one drawer with door for filing, |  |  |  |
|  |  | height/width - 400mm/ 400mm±10%, the other with locker, |  |  |  |
|  |  | height/width- 200mm/ 400mm±10% with Nylonball bearing Sliders. |  |  |  |
|  |  | Extra surface (desk) on main station, |  |  |  |
|  |  | height/length/width- 250mm/1400mm/250mm±10%. |  |  |  |
|  |  | Downward-projecting lamp fixed on extra desk with accessible switch. |  |  |  |
| 28 | Քույրական պոստի սեղան#2 // Nurse station#2 | Plastic laminated consisted of3 wings. | 6 |  |  |
|  |  | Sizes: length/width/height- 1100mm/5000mm/700mm |  |  |  |
|  |  | Station should have 2 drawers, one drawer with door for filing, |  |  |  |
|  |  | height/width - 400mm/ 400mm±10%, the other with locker, |  |  |  |
|  |  | height/width- 200mm/ 400mm±10% with Nylonball bearing Sliders. |  |  |  |
|  |  | Extra surface (desk) on main station, |  |  |  |
| 29 | Բժշկական գործիքների սեղան՝ տեսակ Բ // Table for medical instruments (B) | Table Designed for Surgical Instruments Preparation and Display | 15 |  |  |
|  |  | Shall be made entirely of Stainless Steel , Top Tray shall have Rounded-Edges andshall be Sound-Deadened |  |  |  |
|  |  | Heavy Gauge Stainless Steel Framing, Mobile on Anti-Static Casters, Swiveling with Brakes and mounted on Noiseless Ball BearingsGuard Rails shall be as Optional |  |  |  |
|  |  | Indicative Dimensions: 1400/600mm |  |  |  |
|  |  | Height: from 700 to 800mm |  |  |  |
| 30 | Բժշկական գործիքների սեղան՝ տեսակ Ա // Table for medical instruments (A) | Table Designed for Surgical Instruments Preparation and Display | 35 |  |  |
|  |  | Shall be made entirely of Stainless Steel , Top Tray shall have Rounded-Edges andshall be Sound-Deadened |  |  |  |
|  |  | Heavy Gauge Stainless Steel Framing, Mobile on Anti-Static Casters, Swiveling with Brakes and mounted on Noiseless Ball BearingsGuard Rails shall be as Optional |  |  |  |
|  |  | Indicative Dimensions: 500/600mm |  |  |  |
|  |  | Height: from 700 to 800mm. |  |  |  |
| 31 | Աղբարկղ // Waste bin | Constructed from stainless steel or chrome-plated steel | 190 |  |  |
|  |  | Capacity: not less 12L |  |  |  |
|  |  | Washable surface, tightly closing cover. Pedal action |  |  |  |
| 32 | Շիրմա // Folding screen | Washable. Metallic carcass. With three moveable leaves. | 23 |  |  |
|  |  | Dimensions: |  |  |  |
|  |  | Height: 1800 mm, deviation not more than 10% |  |  |  |
|  |  | Length: from 1700 to 1800 mm |  |  |  |
|  |  | Length of leaf: from 500 to 600 mm |  |  |  |
| 33 | Բժշկական աստիճան // Mounting step | Constructed from chrome-plated or stainless steel tube. Q-ty of steps - 2 | 12 |  |  |
| 34 | Շտատիվ // Infusion fluid holder | IV Pole , Mobile on Castors | 70 |  |  |
|  |  | Adjustable height |  |  |  |
|  |  | 2 Hooks |  |  |  |
| 35 | Պտտվող աթոռ գազային պոմպով // Revolving chair with gas pump | Surgeon’s chair with 5 casters. | 12 |  |  |
|  |  | Variable height: from 50 to 60 cm |  |  |  |
|  |  | Upholstered seat, w/gas pump |  |  |  |
| 36 | Մեկ դռնանի հանդերձարանի պահարան // Changing room locker, single door | Changing room locker with low level shelf with height of 300mm and 2 coat hooks. | 70 |  |  |
|  |  | Washable finish. Ventilated. Nameplate on door. |  |  |  |
|  |  | Lockable, with two keys and key tags with printable label. |  |  |  |
|  |  | Dimensions: |  |  |  |
|  |  | Width: 300mm, deviation not more than 10% |  |  |  |
|  |  | Depth: 450mm, deviation not more than 10% |  |  |  |
|  |  | Height: 1700mm, deviation not more than 10% |  |  |  |
| 37 | Սառնարան խոհանոցային // Kitchen refrigerator | Temperature controlled within the range of from +2° to +8° C. | 12 |  |  |
|  |  | Volume not less 250L |  |  |  |
|  |  | Number of shelves: 3 or 4. |  |  |  |
|  |  | With defrosting system |  |  |  |
|  |  | Power: 220V/50 Hz. deviation not more than 10% |  |  |  |
| 38 | Խոհանոցային կահույք (խոհանոցային լվացարան էլեկտրական սալիկով)#1 // Kitchen sink with electric hod#1 | Dimensions: | 1 |  |  |
|  |  | LxWxH: 1250mmx3500mmx900mm, deviation not more than 10% |  |  |  |
|  |  | Should have ցeramic electric hob |  |  |  |
|  |  | Cupboard with two cabins, each with two-leaved. |  |  |  |
|  |  | Laminate. Table side should be made from plastic waterproof board. |  |  |  |
|  |  | The Purchaser could ask the changes of the dimensions during the contract implementation, subject such changes will not create some additional costs for the Supplier. |  |  |  |
|  |  | Manufactured from chrome plated steel tube |  |  |  |
| 39 | Խոհանոցային կահույք (խոհանոցային լվացարան էլեկտրական սալիկով)#2// Kitchen sink with electric hod#2 | Dimensions: | 1 |  |  |
|  |  | LxWxH: 1250mmx4000mmx900mm, deviation not more than 10% |  |  |  |
|  |  | Should have ցeramic electric hob |  |  |  |
|  |  | Cupboard with two cabins, each with two-leaved. |  |  |  |
|  |  | Laminate. Table side should be made from plastic waterproof board. |  |  |  |
|  |  | The Purchaser could ask the changes of the dimensions during the contract implementation, subject such changes will not create some additional costs for the Supplier. |  |  |  |
|  |  | Manufactured from chrome plated steel tube |  |  |  |
| 40 | Խոհանոցային կահույք (խոհանոցային լվացարան էլեկտրական սալիկով)#3// Kitchen sink with electric hod#3 | Dimensions: | 4 |  |  |
|  |  | LxWxH: 1250mmx4500mmx900mm, deviation not more than 10% |  |  |  |
|  |  | Should have ցeramic electric hob |  |  |  |
|  |  | Cupboard with two cabins, each with two-leaved. |  |  |  |
|  |  | Laminate. Table side should be made from plastic waterproof board. |  |  |  |
|  |  | The Purchaser could ask the changes of the dimensions during the contract implementation, subject such changes will not create some additional costs for the Supplier. |  |  |  |
|  |  | Manufactured from chrome plated steel tube |  |  |  |
| 41 | Պացիենտի սայլակ // Patient trolley | Manufactured from chrome plated steel tube | 6 |  |  |
|  |  | Upholstered fixed Top covered with washable plastic material , Flame Retardant |  |  |  |
|  |  | Complete with side Rails and IV Stand |  |  |  |
|  |  | Mounted on 4 casters 100 mm two of which with brakes |  |  |  |
|  |  | Knock down construction. |  |  |  |
|  |  | Length min: 2000 mm, deviation not more than 10% |  |  |  |
|  |  | Height: 700 mm, deviation not more than 10% |  |  |  |
| 42 | Լվացքի մեքենա // Washing machine - | Features & Specification Features: 7kg load capacity 1400rpm spin speed Energy rating: a Wash rating: a Spin rating: B, | 5 |  |  |
|  |  | not less than 13 wash programmers |  |  |  |
|  |  | Programmed progress indication Delay timer Mini load Reduce creases feature Extra rinse – for sensitive skin Time 4 You and sport wash options Wool programme Warranty: 1 Year |  |  |  |
|  |  | Manufacturer's Guarantee Dimensions: Height: max.850mm Dimensions (WxDxH) 60 cm x 60 cm x 85 cm , Colour White , Washing Efficiency Class A, Spin Class Class B, Energy Class Class A , Maximum Loading Capacity not less 7 kg, Loading Type Front loading , |  |  |  |
|  |  | Maximum Spin Speed 1400 rpm, Average Water Consumption per Cycle 62 liter |  |  |  |
|  |  | Energy Consumption per Cycle max.1.4 kWh |  |  |  |
|  |  | Power requirement 220V/50 Hz |  |  |  |

4. Drawings

These Bidding Documents includes “the following” drawings.

| List of Drawings |  |  |
| --- | --- | --- |
| Drawing Nr. | Drawing Name | Purpose |
| 1 | 1st floor of “Vanadzor MC” CJSC | Installation planning |
| 2 | 3rd floor of “Vanadzor MC” CJSC | Installation planning |
| 3 | Room for installation of Haemodialysis Machine and accessories | Installation planning |
| 4 | Room for installation of Haemodialysis Machine and accessories | Installation planning |

5. Inspections and Tests

The following inspections and tests shall be performed:

For Lots 1; 2 and 3

- The Supplier should conduct the following tests and verification supplied equipment:

- Test ensuring that there are no physical damages to supplied, goods before delivery to the final destination. The Manufacturer’s factory testing certificate for goods supplied will be provided with delivery documents, according to SCC (in case of supply based on CIP terms).

B. The following tests and inspections shall be carried out by the Supplier (or its designated representative) to the Final destination and conducted in the presence of the Purchaser, Beneficiary or their designated representative.

- Verify the completeness of consumables and accessories and other spare parts (if applicable) of the equipments supplied.

- Testing the operation of main function of the equipments supplied, after installation.

C. The Supplier shall provide the Purchaser with a report of the results of above mentioned tests and with a report of completion of training. The names of the trained personal shall be detailed in the report. After that an appropriate delivery-acceptance act shall be issued and signed by the Purchaser, Beneficiary and Supplier.

For Lot- 4

A. The Supplier should conduct the following tests and verification supplied equipment:

- Test ensuring that there are no physical damages to supplied, goods (Medical Furniture), before delivery to the final destination. The Manufacturer’s factory testing certificate for goods supplied will be provided with delivery documents, according to SCC (in case of supply based on CIP terms).

B. The following tests and inspections shall be carried out by the Supplier (or its designated representative) to the Final destination and conducted in the presence of the Purchaser, Beneficiary or their designated representative.

- Verify the completeness of accessories and other spare parts (if applicable) of the Medical Furniture supplied.

- Testing the operation of main function of the Medical Furniture supplied, after installation.

C. The Supplier shall provide the Purchaser with a report of the results of above mentioned tests.

After that an appropriate delivery-acceptance act shall be issued and signed by the Purchaser, Beneficiary and Supplier.

Beneficiary: is director or authorized representative from "Vanadzor MC" CJSC.

# PART 3 - Contract

Section VIII. General Conditions of Contract

Table of Clauses

1.	Definitions	189

2.	Contract Documents	190

3.	Fraud and Corruption	190

4.	Interpretation	190

5.	Language	191

6.	Joint Venture, Consortium or Association	192

7.	Eligibility	192

8.	Notices	192

9.	Governing Law	192

10.	Settlement of Disputes	192

11.	Inspections and Audit by the Bank	193

12.	Scope of Supply	194

13.	Delivery and Documents	194

14.	Supplier’s Responsibilities	194

15.	Contract Price	194

16.	Terms of Payment	194

17.	Taxes and Duties	194

18.	Performance Security	195

19.	Copyright	195

20.	Confidential Information	195

21.	Subcontracting	196

22.	Specifications and Standards	197

23.	Packing and Documents	197

24.	Insurance	197

25.	Transportation and Incidental Services	198

26.	Inspections and Tests	198

27.	Liquidated Damages	199

28.	Warranty	200

29.	Patent Indemnity	200

30.	Limitation of Liability	202

31.	Change in Laws and Regulations	202

32.	Force Majeure	202

33.	Change Orders and Contract Amendments	203

34.	Extensions of Time	203

35.	Termination	204

36.	Assignment	205

37.	Export Restriction	205

Section VIII.  General Conditions of Contract

| 1. Definitions |  | 1.1 The following words and expressions shall have the meanings hereby assigned to them: ### “Bank” means the World Bank and refers to the International Bank for Reconstruction and Development (IBRD) or the International Development Association (IDA). ### “Contract” means the Contract Agreement entered into between the Purchaser and the Supplier, together with the Contract Documents referred to therein, including all attachments, appendices, and all documents incorporated by reference therein. ### “Contract Documents” means the documents listed in the Contract Agreement, including any amendments thereto. ### “Contract Price” means the price payable to the Supplier as specified in the Contract Agreement, subject to such additions and adjustments thereto or deductions therefrom, as may be made pursuant to the Contract. ### “Day” means calendar day. ### “Completion” means the fulfillment of the Related Services by the Supplier in accordance with the terms and conditions set forth in the Contract. ### “GCC” means the General Conditions of Contract. ### “Goods” means all of the commodities, raw material, machinery and equipment, and/or other materials that the Supplier is required to supply to the Purchaser under the Contract. ### “Purchaser’s Country” is the country specified in the Special Conditions of Contract (SCC). ### “Purchaser” means the entity purchasing the Goods and Related Services, as specified in the SCC. ### “Related Services” means the services incidental to the supply of the goods, such as insurance, installation, training and initial maintenance and other such obligations of the Supplier under the Contract. ### “SCC” means the Special Conditions of Contract. ### “Subcontractor” means any person, private or government entity, or a combination of the above, to whom any part of the Goods to be supplied or execution of any part of the Related Services is subcontracted by the Supplier. ### “Supplier” means the person, private or government entity, or a combination of the above, whose bid to perform the Contract has been accepted by the Purchaser and is named as such in the Contract Agreement. ### “The Project Site,” where applicable, means the place named in the SCC. |  |
| --- | --- | --- | --- |
| 2. Contract Documents |  | - Subject to the order of precedence set forth in the Contract Agreement, all documents forming the Contract (and all parts thereof) are intended to be correlative, complementary, and mutually explanatory. The Contract Agreement shall be read as a whole. |  |
| 3. Corrupt and Fraudulent Practices |  | 3.1 The Bank requires compliance with its policy in regard to corrupt and fraudulent practices as set forth in Appendix to the GCC. 3.2 The Purchaser requires the Supplier to disclose any commissions or fees that may have been paid or are to be paid to agents or any other party with respect to the bidding process or execution of the Contract. The information disclosed must include at least the name and address of the agent or other party, the amount and currency, and the purpose of the commission, gratuity or fee. |  |
| 4. Interpretation |  | - If the context so requires it, singular means plural and vice versa. - Incoterms ### Unless inconsistent with any provision of the Contract, the meaning of any trade term and the rights and obligations of parties thereunder shall be as prescribed by Incoterms. ### The terms EXW, CIP, FCA, CFR and other similar terms, when used, shall be governed by the rules prescribed in the current edition of Incoterms specified in the SCC and published by the International Chamber of Commerce in Paris, France. - Entire Agreement The Contract constitutes the entire agreement between the Purchaser and the Supplier and supersedes all communications, negotiations and agreements (whether written or oral) of the parties with respect thereto made prior to the date of Contract. - Amendment No amendment or other variation of the Contract shall be valid unless it is in writing, is dated, expressly refers to the Contract, and is signed by a duly authorized representative of each party thereto. - Nonwaiver ### Subject to GCC Sub-Clause 4.5(b) below, no relaxation, forbearance, delay, or indulgence by either party in enforcing any of the terms and conditions of the Contract or the granting of time by either party to the other shall prejudice, affect, or restrict the rights of that party under the Contract, neither shall any waiver by either party of any breach of Contract operate as waiver of any subsequent or continuing breach of Contract. ### Any waiver of a party’s rights, powers, or remedies under the Contract must be in writing, dated, and signed by an authorized representative of the party granting such waiver, and must specify the right and the extent to which it is being waived. - Severability If any provision or condition of the Contract is prohibited or rendered invalid or unenforceable, such prohibition, invalidity or unenforceability shall not affect the validity or enforceability of any other provisions and conditions of the Contract. |  |
| 5. Language |  | - The Contract as well as all correspondence and documents relating to the Contract exchanged by the Supplier and the Purchaser, shall be written in the language specified in the SCC. Supporting documents and printed literature that are part of the Contract may be in another language provided they are accompanied by an accurate translation of the relevant passages in the language specified, in which case, for purposes of interpretation of the Contract, this translation shall govern. - The Supplier shall bear all costs of translation to the governing language and all risks of the accuracy of such translation, for documents provided by the Supplier. |  |
| 6. Joint Venture, Consortium or Association |  | - If the Supplier is a joint venture, consortium, or association, all of the parties shall be jointly and severally liable to the Purchaser for the fulfillment of the provisions of the Contract and shall designate one party to act as a leader with authority to bind the joint venture, consortium, or association. The composition or the constitution of the joint venture, consortium, or association shall not be altered without the prior consent of the Purchaser. |  |
| 7. Eligibility |  | - The Supplier and its Subcontractors shall have the nationality of an eligible country. A Supplier or Subcontractor shall be deemed to have the nationality of a country if it is a citizen or constituted, incorporated, or registered, and operates in conformity with the provisions of the laws of that country. - All Goods and Related Services to be supplied under the Contract and financed by the Bank shall have their origin in Eligible Countries. For the purpose of this Clause, origin means the country where the goods have been grown, mined, cultivated, produced, manufactured, or processed; or through manufacture, processing, or assembly, another commercially recognized article results that differs substantially in its basic characteristics from its components. |  |
| 8. Notices |  | - Any notice given by one party to the other pursuant to the Contract shall be in writing to the address specified in the SCC. The term “in writing” means communicated in written form with proof of receipt. - A notice shall be effective when delivered or on the notice’s effective date, whichever is later. |  |
| 9. Governing Law | - The Contract shall be governed by and interpreted in accordance with the laws of the Purchaser’s Country, unless otherwise specified in the SCC. - Throughout the execution of the Contract, the Contractor shall comply with the import of goods and services prohibitions in the Purchaser’s country when (a) as a matter of law or official regulations, the Borrower’s country prohibits commercial relations with that country; or - (b) by an act of compliance with a decision of the United Nations Security Council taken under Chapter VII of the Charter of the United Nations, the Borrower’s Country prohibits any import of goods from that country or any payments to any country, person, or entity in that country. |
| 10 Settlement of Disputes | - The Purchaser and the Supplier shall make every effort to resolve amicably by direct informal negotiation any disagreement or dispute arising between them under or in connection with the Contract. - If, after twenty-eight (28) days, the parties have failed to resolve their dispute or difference by such mutual consultation, then either the Purchaser or the Supplier may give notice to the other party of its intention to commence arbitration, as hereinafter provided, as to the matter in dispute, and no arbitration in respect of this matter may be commenced unless such notice is given. Any dispute or difference in respect of which a notice of intention to commence arbitration has been given in accordance with this Clause shall be finally settled by arbitration. Arbitration may be commenced prior to or after delivery of the Goods under the Contract. Arbitration proceedings shall be conducted in accordance with the rules of procedure specified in the SCC. - Notwithstanding any reference to arbitration herein, - the parties shall continue to perform their respective obligations under the Contract unless they otherwise agree; and - the Purchaser shall pay the Supplier any monies due the Supplier. |
| 11. Inspections and Audit by the Bank | - The Supplier shall keep, and shall make all reasonable efforts to cause its Subcontractors to keep, accurate and systematic accounts and records in respect of the Goods in such form and details as will clearly identify relevant time changes and costs. - The Supplier shall permit, and shall cause its Subcontractorsto permit, the Bank and/or persons appointed by the Bank to inspect the Supplier’s offices and all accounts and records relating to the performance of the Contract and the submission of the bid, and to have such accounts and records audited by auditors appointed by the Bank if requested by the Bank. The Supplier’s and its Subcontractors and consultants’ attention is drawn to Clause 3 [Fraud and Corruption], which provides, inter alia, that acts intended to materially impede the exercise of the Bank’s inspection and audit rights provided for under this Sub-Clause 11.1 constitute a prohibited practice subject to contract termination (as well as to a determination of ineligibility pursuant to the Bank’s prevailing sanctions procedures) |
| 12. Scope of Supply | 12.1 The Goods and Related Services to be supplied shall be as specified in the Schedule of Requirements. |
| 13. Delivery and Documents | 13.1 Subject to GCC Sub-Clause 33.1, the Delivery of the Goods and Completion of the Related Services shall be in accordance with the Delivery and Completion Schedule specified in the Schedule of Requirements. The details of shipping and other documents to be furnished by the Supplier are specified in the SCC. |
| 14. Supplier’s Responsibilities | 14.1 The Supplier shall supply all the Goods and Related Services included in the Scope of Supply in accordance with GCC Clause 12, and the Delivery and Completion Schedule, as per GCC Clause 13. |
| 15 Contract Price | 15.1 Prices charged by the Supplier for the Goods supplied and the Related Services performed under the Contract shall not vary from the prices quoted by the Supplier in its bid, with the exception of any price adjustments authorized in the SCC. |
| 16. Terms of Payment | 16.1 The Contract Price, including any Advance Payments, if applicable, shall be paid as specified in the SCC. 16.2 The Supplier’s request for payment shall be made to the Purchaser in writing, accompanied by invoices describing, as appropriate, the Goods delivered and Related Services performed, and by the documents submitted pursuant to GCC Clause 13 and upon fulfillment of all other obligations stipulated in the Contract. 16.3 Payments shall be made promptly by the Purchaser, but in no case later than sixty (60) days after submission of an invoice or request for payment by the Supplier, and after the Purchaser has accepted it. 16.4 The currencies in which payments shall be made to the Supplier under this Contract shall be those in which the bid price is expressed. 16.5 In the event that the Purchaser fails to pay the Supplier any payment by its due date or within the period set forth in the SCC, the Purchaser shall pay to the Supplier interest on the amount of such delayed payment at the rate shown in the SCC, for the period of delay until payment has been made in full, whether before or after judgment or arbitrage award. |
| 17. Taxes and Duties | 17.1 For goods manufactured outside the Purchaser’s Country, the Supplier shall be entirely responsible for all taxes, stamp duties, license fees, and other such levies imposed outside the Purchaser’s Country. 17.2 For goods Manufactured within the Purchaser’s country, the Supplier shall be entirely responsible for all taxes, duties, license fees, etc., incurred until delivery of the contracted Goods to the Purchaser. 17.3 If any tax exemptions, reductions, allowances or privileges may be available to the Supplier in the Purchaser’s Country, the Purchaser shall use its best efforts to enable the Supplier to benefit from any such tax savings to the maximum allowable extent. |
| 18. Performance Security | 18.1 If required as specified in the SCC, the Supplier shall, within twenty-eight (28) days of the notification of contract award, provide a performance security for the performance of the Contract in the amount specified in the SCC. 18.2 The proceeds of the Performance Security shall be payable to the Purchaser as compensation for any loss resulting from the Supplier’s failure to complete its obligations under the Contract. 18.3 As specified in the SCC, the Performance Security, if required, shall be denominated in the currency(ies) of the Contract, or in a freely convertible currency acceptable to the Purchaser; and shall be in one of the format stipulated by the Purchaser in the SCC, or in another format acceptable to the Purchaser. 18.4 The Performance Security shall be discharged by the Purchaser and returned to the Supplier not later than twenty-eight (28) days following the date of Completion of the Supplier’s performance obligations under the Contract, including any warranty obligations, unless specified otherwise in the SCC. |
| 19. Copyright | 19.1 The copyright in all drawings, documents, and other materials containing data and information furnished to the Purchaser by the Supplier herein shall remain vested in the Supplier, or, if they are furnished to the Purchaser directly or through the Supplier by any third party, including suppliers of materials, the copyright in such materials shall remain vested in such third party |
| 20. Confidential Information | 20.1 The Purchaser and the Supplier shall keep confidential and shall not, without the written consent of the other party hereto, divulge to any third party any documents, data, or other information furnished directly or indirectly by the other party hereto in connection with the Contract, whether such information has been furnished prior to, during or following completion or termination of the Contract. Notwithstanding the above, the Supplier may furnish to its Subcontractor such documents, data, and other information it receives from the Purchaser to the extent required for the Subcontractor to perform its work under the Contract, in which event the Supplier shall obtain from such Subcontractor an undertaking of confidentiality similar to that imposed on the Supplier under GCC Clause 20. 20.2 The Purchaser shall not use such documents, data, and other information received from the Supplier for any purposes unrelated to the contract. Similarly, the Supplier shall not use such documents, data, and other information received from the Purchaser for any purpose other than the performance of the Contract. 20.3 The obligation of a party under GCC Sub-Clauses 20.1 and 20.2 above, however, shall not apply to information that: ### the Purchaser or Supplier need to share with the Bank or other institutions participating in the financing of the Contract; ### now or hereafter enters the public domain through no fault of that party; ### can be proven to have been possessed by that party at the time of disclosure and which was not previously obtained, directly or indirectly, from the other party; or ### otherwise lawfully becomes available to that party from a third party that has no obligation of confidentiality. 20.4 The above provisions of GCC Clause 20 shall not in any way modify any undertaking of confidentiality given by either of the parties hereto prior to the date of the Contract in respect of the Supply or any part thereof. 20.5 The provisions of GCC Clause 20 shall survive completion or termination, for whatever reason, of the Contract. |
| 21. Subcontracting | 21.1 The Supplier shall notify the Purchaser in writing of all subcontracts awarded under the Contract if not already specified in the bid. Such notification, in the original bid or later shall not relieve the Supplier from any of its obligations, duties, responsibilities, or liability under the Contract. 21.2 Subcontracts shall comply with the provisions of GCC Clauses 3 and 7. |
| 22. Specifications and Standards | 22.1 Technical Specifications and Drawings ### The Goods and Related Services supplied under this Contract shall conform to the technical specifications and standards mentioned in Section VI, Schedule of Requirements and, when no applicable standard is mentioned, the standard shall be equivalent or superior to the official standards whose application is appropriate to the Goods’ country of origin. ### The Supplier shall be entitled to disclaim responsibility for any design, data, drawing, specification or other document, or any modification thereof provided or designed by or on behalf of the Purchaser, by giving a notice of such disclaimer to the Purchaser. ### Wherever references are made in the Contract to codes and standards in accordance with which it shall be executed, the edition or the revised version of such codes and standards shall be those specified in the Schedule of Requirements. During Contract execution, any changes in any such codes and standards shall be applied only after approval by the Purchaser and shall be treated in accordance with GCC Clause 33. |
| 23. Packing and Documents | 23.1 The Supplier shall provide such packing of the Goods as is required to prevent their damage or deterioration during transit to their final destination, as indicated in the Contract. During transit, the packing shall be sufficient to withstand, without limitation, rough handling and exposure to extreme temperatures, salt and precipitation, and open storage. Packing case size and weights shall take into consideration, where appropriate, the remoteness of the goods’ final destination and the absence of heavy handling facilities at all points in transit. 23.2 The packing, marking, and documentation within and outside the packages shall comply strictly with such special requirements as shall be expressly provided for in the Contract, including additional requirements, if any, specified in the SCC, and in any other instructions ordered by the Purchaser. |
| 24. Insurance | 24.1 Unless otherwise specified in the SCC, the Goods supplied under the Contract shall be fully insured—in a freely convertible currency from an eligible country—against loss or damage incidental to manufacture or acquisition, transportation, storage, and delivery, in accordance with the applicable Incoterms or in the manner specified in the SCC. |
| 25. Transportation and Incidental Services | 25.1 Unless otherwise specified in the SCC, responsibility for arranging transportation of the Goods shall be in accordance with the specified Incoterms. |
|  | 25.2 The Supplier may be required to provide any or all of the following services, including additional services, if any, specified in SCC: (a) performance or supervision of on-site assembly and/or start-up of the supplied Goods; (b) furnishing of tools required for assembly and/or maintenance of the supplied Goods; (c) furnishing of a detailed operations and maintenance manual for each appropriate unit of the supplied Goods; (d) performance or supervision or maintenance and/or repair of the supplied Goods, for a period of time agreed by the parties, provided that this service shall not relieve the Supplier of any warranty obligations under this Contract; and (e) training of the Purchaser’s personnel, at the Supplier’s plant and/or on-site, in assembly, start-up, operation, maintenance, and/or repair of the supplied Goods. 25.3 Prices charged by the Supplier for incidental services, if not included in the Contract Price for the Goods, shall be agreed upon in advance by the parties and shall not exceed the prevailing rates charged to other parties by the Supplier for similar services |
| 26. Inspections and Tests | 26.1 The Supplier shall at its own expense and at no cost to the Purchaser carry out all such tests and/or inspections of the Goods and Related Services as are specified in the SCC. 26.2 The inspections and tests may be conducted on the premises of the Supplier or its Subcontractor, at point of delivery, and/or at the Goods’ final destination, or in another place in the Purchaser’s Country as specified in the SCC. Subject to GCC Sub-Clause 26.3, if conducted on the premises of the Supplier or its Subcontractor, all reasonable facilities and assistance, including access to drawings and production data, shall be furnished to the inspectors at no charge to the Purchaser. 26.3 The Purchaser or its designated representative shall be entitled to attend the tests and/or inspections referred to in GCC Sub-Clause 26.2, provided that the Purchaser bear all of its own costs and expenses incurred in connection with such attendance including, but not limited to, all traveling and board and lodging expenses. 26.4 Whenever the Supplier is ready to carry out any such test and inspection, it shall give a reasonable advance notice, including the place and time, to the Purchaser. The Supplier shall obtain from any relevant third party or manufacturer any necessary permission or consent to enable the Purchaser or its designated representative to attend the test and/or inspection. 26.5 The Purchaser may require the Supplier to carry out any test and/or inspection not required by the Contract but deemed necessary to verify that the characteristics and performance of the Goods comply with the technical specifications codes and standards under the Contract, provided that the Supplier’s reasonable costs and expenses incurred in the carrying out of such test and/or inspection shall be added to the Contract Price. Further, if such test and/or inspection impedes the progress of manufacturing and/or the Supplier’s performance of its other obligations under the Contract, due allowance will be made in respect of the Delivery Dates and Completion Dates and the other obligations so affected. 26.6 The Supplier shall provide the Purchaser with a report of the results of any such test and/or inspection. 26.7 The Purchaser may reject any Goods or any part thereof that fail to pass any test and/or inspection or do not conform to the specifications. The Supplier shall either rectify or replace such rejected Goods or parts thereof or make alterations necessary to meet the specifications at no cost to the Purchaser, and shall repeat the test and/or inspection, at no cost to the Purchaser, upon giving a notice pursuant to GCC Sub-Clause 26.4. 26.8 The Supplier agrees that neither the execution of a test and/or inspection of the Goods or any part thereof, nor the attendance by the Purchaser or its representative, nor the issue of any report pursuant to GCC Sub-Clause 26.6, shall release the Supplier from any warranties or other obligations under the Contract. |
| 27. Liquidated Damages | 27.1 Except as provided under GCC Clause 32, if the Supplier fails to deliver any or all of the Goods by the Date(s) of delivery or perform the Related Services within the period specified in the Contract, the Purchaser may without prejudice to all its other remedies under the Contract, deduct from the Contract Price, as liquidated damages, a sum equivalent to the percentage specified in the SCC of the delivered price of the delayed Goods or unperformed Services for each week or part thereof of delay until actual delivery or performance, up to a maximum deduction of the percentage specified in those SCC. Once the maximum is reached, the Purchaser may terminate the Contract pursuant to GCC Clause 35. |
| 28. Warranty | 28.1 The Supplier warrants that all the Goods are new, unused, and of the most recent or current models, and that they incorporate all recent improvements in design and materials, unless provided otherwise in the Contract. 28.2 Subject to GCC Sub-Clause 22.1(b), the Supplier further warrants that the Goods shall be free from defects arising from any act or omission of the Supplier or arising from design, materials, and workmanship, under normal use in the conditions prevailing in the country of final destination. 28.3 Unless otherwise specified in the SCC, the warranty shall remain valid for twelve (12) months after the Goods, or any portion thereof as the case may be, have been delivered to and accepted at the final destination indicated in the SCC, or for eighteen (18) months after the date of shipment from the port or place of loading in the country of origin, whichever period concludes earlier. 28.4 The Purchaser shall give notice to the Supplier stating the nature of any such defects together with all available evidence thereof, promptly following the discovery thereof. The Purchaser shall afford all reasonable opportunity for the Supplier to inspect such defects. 28.5 Upon receipt of such notice, the Supplier shall, within the period specified in the SCC, expeditiously repair or replace the defective Goods or parts thereof, at no cost to the Purchaser. 28.6 If having been notified, the Supplier fails to remedy the defect within the period specified in the SCC, the Purchaser may proceed to take within a reasonable period such remedial action as may be necessary, at the Supplier’s risk and expense and without prejudice to any other rights which the Purchaser may have against the Supplier under the Contract. |
| 29. Patent Indemnity | 29.1 The Supplier shall, subject to the Purchaser’s compliance with GCC Sub-Clause 29.2, indemnify and hold harmless the Purchaser and its employees and officers from and against any and all suits, actions or administrative proceedings, claims, demands, losses, damages, costs, and expenses of any nature, including attorney’s fees and expenses, which the Purchaser may suffer as a result of any infringement or alleged infringement of any patent, utility model, registered design, trademark, copyright, or other intellectual property right registered or otherwise existing at the date of the Contract by reason of: ### the installation of the Goods by the Supplier or the use of the Goods in the country where the Site is located; and ### the sale in any country of the products produced by the Goods. ### Such indemnity shall not cover any use of the Goods or any part thereof other than for the purpose indicated by or to be reasonably inferred from the Contract, neither any infringement resulting from the use of the Goods or any part thereof, or any products produced thereby in association or combination with any other equipment, plant, or materials not supplied by the Supplier, pursuant to the Contract. 29.2 If any proceedings are brought or any claim is made against the Purchaser arising out of the matters referred to in GCC Sub-Clause 29.1, the Purchaser shall promptly give the Supplier a notice thereof, and the Supplier may at its own expense and in the Purchaser’s name conduct such proceedings or claim and any negotiations for the settlement of any such proceedings or claim. 29.3 If the Supplier fails to notify the Purchaser within twenty-eight (28) days after receipt of such notice that it intends to conduct any such proceedings or claim, then the Purchaser shall be free to conduct the same on its own behalf. 29.4 The Purchaser shall, at the Supplier’s request, afford all available assistance to the Supplier in conducting such proceedings or claim, and shall be reimbursed by the Supplier for all reasonable expenses incurred in so doing. 29.5 The Purchaser shall indemnify and hold harmless the Supplier and its employees, officers, and Subcontractors from and against any and all suits, actions or administrative proceedings, claims, demands, losses, damages, costs, and expenses of any nature, including attorney’s fees and expenses, which the Supplier may suffer as a result of any infringement or alleged infringement of any patent, utility model, registered design, trademark, copyright, or other intellectual property right registered or otherwise existing at the date of the Contract arising out of or in connection with any design, data, drawing, specification, or other documents or materials provided or designed by or on behalf of the Purchaser. |
| 30 Limitation of Liability | 30.1 Except in cases of criminal negligence or willful misconduct, (a) the Supplier shall not be liable to the Purchaser, whether in contract, tort, or otherwise, for any indirect or consequential loss or damage, loss of use, loss of production, or loss of profits or interest costs, provided that this exclusion shall not apply to any obligation of the Supplier to pay liquidated damages to the Purchaser and (b) the aggregate liability of the Supplier to the Purchaser, whether under the Contract, in tort or otherwise, shall not exceed the total Contract Price, provided that this limitation shall not apply to the cost of repairing or replacing defective equipment, or to any obligation of the supplier to indemnify the purchaser with respect to patent infringement |
| 31. Change in Laws and Regulations | 31.1 Unless otherwise specified in the Contract, if after the date of 28 days prior to date of Bid submission, any law, regulation, ordinance, order or bylaw having the force of law is enacted, promulgated, abrogated, or changed in the place of the Purchaser’s country where the Site is located (which shall be deemed to include any change in interpretation or application by the competent authorities) that subsequently affects the Delivery Date and/or the Contract Price, then such Delivery Date and/or Contract Price shall be correspondingly increased or decreased, to the extent that the Supplier has thereby been affected in the performance of any of its obligations under the Contract. Notwithstanding the foregoing, such additional or reduced cost shall not be separately paid or credited if the same has already been accounted for in the price adjustment provisions where applicable, in accordance with GCC Clause 15. |
| 32. Force Majeure | 32.1 The Supplier shall not be liable for forfeiture of its Performance Security, liquidated damages, or termination for default if and to the extent that its delay in performance or other failure to perform its obligations under the Contract is the result of an event of Force Majeure. 32.2 For purposes of this Clause, “Force Majeure” means an event or situation beyond the control of the Supplier that is not foreseeable, is unavoidable, and its origin is not due to negligence or lack of care on the part of the Supplier. Such events may include, but not be limited to, acts of the Purchaser in its sovereign capacity, wars or revolutions, fires, floods, epidemics, quarantine restrictions, and freight embargoes. 32.3 If a Force Majeure situation arises, the Supplier shall promptly notify the Purchaser in writing of such condition and the cause thereof. Unless otherwise directed by the Purchaser in writing, the Supplier shall continue to perform its obligations under the Contract as far as is reasonably practical, and shall seek all reasonable alternative means for performance not prevented by the Force Majeure event. |
| 33. Change Orders and Contract Amendments | 33.1 The Purchaser may at any time order the Supplier through notice in accordance GCC Clause 8, to make changes within the general scope of the Contract in any one or more of the following: ### drawings, designs, or specifications, where Goods to be furnished under the Contract are to be specifically manufactured for the Purchaser; ### the method of shipment or packing; ### the place of delivery; and ### the Related Services to be provided by the Supplier. 33.2 If any such change causes an increase or decrease in the cost of, or the time required for, the Supplier’s performance of any provisions under the Contract, an equitable adjustment shall be made in the Contract Price or in the Delivery/Completion Schedule, or both, and the Contract shall accordingly be amended. Any claims by the Supplier for adjustment under this Clause must be asserted within twenty-eight (28) days from the date of the Supplier’s receipt of the Purchaser’s change order. 33.3 Prices to be charged by the Supplier for any Related Services that might be needed but which were not included in the Contract shall be agreed upon in advance by the parties and shall not exceed the prevailing rates charged to other parties by the Supplier for similar services. 33.4 Subject to the above, no variation in or modification of the terms of the Contract shall be made except by written amendment signed by the parties. |
| 34. Extensions of Time | 34.1 If at any time during performance of the Contract, the Supplier or its subcontractors should encounter conditions impeding timely delivery of the Goods or completion of Related Services pursuant to GCC Clause 13, the Supplier shall promptly notify the Purchaser in writing of the delay, its likely duration, and its cause. As soon as practicable after receipt of the Supplier’s notice, the Purchaser shall evaluate the situation and may at its discretion extend the Supplier’s time for performance, in which case the extension shall be ratified by the parties by amendment of the Contract. 34.2 Except in case of Force Majeure, as provided under GCC Clause 32, a delay by the Supplier in the performance of its Delivery and Completion obligations shall render the Supplier liable to the imposition of liquidated damages pursuant to GCC Clause 26, unless an extension of time is agreed upon, pursuant to GCC Sub-Clause 34.1. |
| 35. Termination | 35.1 Termination for Default ### The Purchaser, without prejudice to any other remedy for breach of Contract, by written notice of default sent to the Supplier, may terminate the Contract in whole or in part: #### if the Supplier fails to deliver any or all of the Goods within the period specified in the Contract, or within any extension thereof granted by the Purchaser pursuant to GCC Clause 34; #### if the Supplier fails to perform any other obligation under the Contract; or #### if the Supplier, in the judgment of the Purchaser has engaged in fraud and corruption, as defined in GCC Clause 3, in competing for or in executing the Contract. ### In the event the Purchaser terminates the Contract in whole or in part, pursuant to GCC Clause 35.1(a), the Purchaser may procure, upon such terms and in such manner as it deems appropriate, Goods or Related Services similar to those undelivered or not performed, and the Supplier shall be liable to the Purchaser for any additional costs for such similar Goods or Related Services. However, the Supplier shall continue performance of the Contract to the extent not terminated. 35.2 Termination for Insolvency. ### The Purchaser may at any time terminate the Contract by giving notice to the Supplier if the Supplier becomes bankrupt or otherwise insolvent. In such event, termination will be without compensation to the Supplier, provided that such termination will not prejudice or affect any right of action or remedy that has accrued or will accrue thereafter to the Purchaser 35.3 Termination for Convenience. ### The Purchaser, by notice sent to the Supplier, may terminate the Contract, in whole or in part, at any time for its convenience. The notice of termination shall specify that termination is for the Purchaser’s convenience, the extent to which performance of the Supplier under the Contract is terminated, and the date upon which such termination becomes effective. ### The Goods that are complete and ready for shipment within twenty-eight (28) days after the Supplier’s receipt of notice of termination shall be accepted by the Purchaser at the Contract terms and prices. For the remaining Goods, the Purchaser may elect: #### to have any portion completed and delivered at the Contract terms and prices; and/or #### to cancel the remainder and pay to the Supplier an agreed amount for partially completed Goods and Related Services and for materials and parts previously procured by the Supplier. |
| 36. Assignment | 36.1 Neither the Purchaser nor the Supplier shall assign, in whole or in part, their obligations under this Contract, except with prior written consent of the other party. |
| 37. Export Restriction | 37.1 Notwithstanding any obligation under the Contract to complete all export formalities, any export restrictions attributable to the Purchaser, to the country of the Purchaser, or to the use of the products/goods, systems or services to be supplied, which arise from trade regulations from a country supplying those products/goods, systems or services, and whichsubstantially impede the Supplier from meeting its obligations under the Contract, shall release the Supplier from the obligation to provide deliveries or services, always provided, however, that the Supplier can demonstrate to the satisfaction of the Purchaser and of the Bank that it has completed all formalities in a timely manner, including applying for permits, authorizations and licenses necessary for the exportof the products/goods, systems or services under the terms of the Contract. Termination of the Contract on this basis shall be for the Purchaser’s convenience pursuant to Sub-Clause 35.3. |

APPENDIX TO GENERAL CONDITIONS

Bank’s Policy- Corrupt and Fraudulent Practices

(text in this Appendix shall not be modified)

Guidelines for Procurement of Goods, Works, and Non-Consulting Services under IBRD Loans and IDA Credits & Grants by World Bank Borrowers, dated January 2011:

“Fraud and Corruption:

1.16	It is the Bank’s policy to require that Borrowers (including beneficiaries of Bank loans), bidders, suppliers, contractors and their agents (whether declared or not), sub-contractors, sub-consultants, service providers or suppliers, and any personnel thereof, observe the highest standard of ethics during the procurement and execution of Bank-financed contracts. In pursuance of this policy, the Bank:

(a)	defines, for the purposes of this provision, the terms set forth below as follows:

(i)	“corrupt practice” is the offering, giving, receiving, or soliciting, directly or indirectly, of anything of value to influence improperly the actions of another party;;

(ii) 	“fraudulent practice” is any act or omission, including a misrepresentation, that knowingly or recklessly misleads, or attempts to mislead, a party to obtain a financial or other benefit or to avoid an obligation;

(iii)	“collusive practice” is an arrangement between two or more parties designed to achieve an improper purpose, including to influence improperly the actions of another party;

(iv)	“coercive practice” is impairing or harming, or threatening to impair or harm, directly or indirectly, any party or the property of the party to influence improperly the actions of a party;

(v)	"obstructive practice" is:

(aa)	deliberately destroying, falsifying, altering, or concealing of evidence material to the investigation or making false statements to investigators in order to materially impede a Bank investigation into allegations of a corrupt, fraudulent, coercive or collusive practice; and/or threatening, harassing or intimidating any party to prevent it from disclosing its knowledge of matters relevant to the investigation or from pursuing the investigation, or

(bb)	acts intended to materially impede the exercise of the Bank’s inspection and audit rights provided for under paragraph 1.16(e) below.

(b)	will reject a proposal for award if it determines that the bidder recommended for award, or any of its personnel, or its agents, or its sub-consultants, sub-contractors, service providers, suppliers and/or their employees, has, directly or indirectly, engaged in corrupt, fraudulent, collusive, coercive, or obstructive practices in competing for the contract in question;

(c)	will declare misprocurement and cancel the portion of the loan allocated to a contract if it determines at any time that representatives of the Borrower or of a recipient of any part of the proceeds of the loan engaged in corrupt, fraudulent, collusive, coercive, or obstructive practices during the procurement or the implementation of the contract in question, without the Borrower having taken timely and appropriate action satisfactory to the Bank to address such practices when they occur, including by failing to inform the Bank in a timely manner at the time they knew of the practices;

(d)	will sanction a firm or individual, at any time, in accordance with the prevailing Bank’s sanctions procedures, including by publicly declaring such firm or individual ineligible, either indefinitely or for a stated period of time: (i) to be awarded a Bank-financed contract; and (ii) to be a nominated;

(e)	will require that a clause be included in bidding documents and in contracts financed by a Bank loan, requiring bidders, suppliers and contractors, and their sub-contractors, agents, personnel, consultants, service providers, or suppliers, to permit the Bank to inspect all accounts, records, and other documents relating to the submission of bids and contract performance, and to have them audited by auditors appointed by the Bank.”

.

| Section IX. Special Conditions of Contract |  |
| --- | --- |
| The following Special Conditions of Contract (SCC) shall supplement and / or amend the General Conditions of Contract (GCC). Whenever there is a conflict, the provisions herein shall prevail over those in the GCC. |  |
| GCC 1.1(i) | The Purchaser’s country is: Republic of Armenia |
| GCC 1.1(j) | The Purchaser is: “Health projects implementation unit” State Agency |
| GCC 1.1 (o) | Final Destination is: Armenia, Lori Marz, Vanadzor Banaki St., 8a Build |
| GCC 4.2 (a) | The meaning of the trade terms shall be as prescribed by Incoterms. |
| GCC 4.2 (b) | The version edition of Incoterms shall be 2010 year |
| GCC 5.1 | The language shall be: English |
| GCC 8.1 | For notices, the Purchaser’s address shall be: Attention: Nelson Zuloyan – Acting Director of the HPIU SA Street Address: Komitas 49/4 avenue Floor/ Room number: 5thfloor City: Yerevan ZIP Code: 0051 Country: Republic of Armenia Telephone: +374 10 297-536; +374 10 297-537 Electronic mail address: info@healthpiu.am |
| GCC 9.1 | The governing law shall be the law of: Republic of Armenia |
| GCC 10.2 | The rules of procedure for arbitration proceedings pursuant to GCC Clause 10.2 shall be as follows: (a)Contract with foreign Supplier: GCC 10.2 (a)—Any dispute, controversy or claim arising out of or relating to this Contract, or breach, termination or invalidity thereof, shall be settled by arbitration in accordance with the UNCITRAL Arbitration Rules as at present in force. (b)Contracts with Supplier national of the Purchaser’s country: In the case of a dispute between the Purchaser and a Supplier who is a national of the Purchaser’s country, the dispute shall be referred to adjudication or arbitration in accordance with the laws of the Purchaser’s country. |
| GCC 13.1 | Details of Shipping and other Documents to be furnished by the Supplier are: A) If goods will be imported: - copies of the Supplier’s invoice showing Goods’ description, quantity, unit price, and total amount; - copies of the packing list identifying contents of each package; - insurance certificate; - Manufacturer’s factory testing certificate for goods supplied (Applicable for Lots 1; 2 and 3); - certificate of origin or equivalent document (for products from EEU countries); - User manuals in English language with Russian/Armenian translation (Not applicable for Item 13 under Lot 2; Items 10, 28, 29, 30 under Lot 3). Required only for Items 22, 23, 24, 37 and 42 under Lot 4. Manufacturer’s or Supplier’s warranty certificate shall be provided after installation of goods. The above documents shall be received by the Purchaser before arrival of the Goods and, if not received, the Supplier will be responsible for any consequent expenses. B) If goods are already imported or manufactured domestically: - copies of the Supplier’s invoice showing Goods’ description, quantity, unit price, and total amount; - certificate of origin or equivalent document (for products from EEU countries); - User manuals in English language with Russian/Armenian translation (Not applicable for Item 13 under Lot 2; Items 10, 28, 29, 30 under Lot 3). Required only for Items 22, 23, 24, 37 and 42 under Lot 4. Manufacturer’s or Supplier’s warranty certificate shall be provided after installation of goods. |
| GCC 15.1 | The prices charged for the Goods supplied and the related Services performed “shall not,” be adjustable. If prices are adjustable, the following method shall be used to calculate the price adjustment: Not applicable |
| GCC 16.1 | Sample provision GCC 16.1—The method and conditions of payment to be made to the Supplier under this Contract shall be as follows: Payment for Goods supplied from abroad: Payment for Goods supplied from abroad shall be made in( ) [currency of the Contract Price] in the following manner: (i) Advance Payment: Ten (10) percent of the Contract Price shall be paid within thirty (30) days of signing of the Contract, and upon submission of claim and a bank guarantee for equivalent amount valid until the Goods are delivered and in the form provided in the bidding documents or another form acceptable to the Purchaser. (ii) On Shipment: Eighty (80) percent of the Contract Price of the Goods shipped shall be paid through irrevocable confirmed letter of credit opened in favor of the Supplier in a bank in its country, upon submission of documents specified in GCC Clause 12. (iii) On Acceptance: Ten (10) percent of the Contract Price of Goods received shall be paid within thirty (30) days of receipt of the Goods upon submission of claim supported by the acceptance certificate issued by the Purchaser. Payment for Provided services shall be made in ______[currency] in the following manner: Ninety (90) percent of the Contract Price of Services received shall be paid within thirty (30) days of presentation of claim supported by a certificate from the Purchaser declaring that the Goods have been delivered and that all other contracted Services have been performed. Payment for Goods and Services supplied from within the Purchaser’s country: Payment for Goods and Services supplied from within the Purchaser’s country shall be made in _____ [currency], as follows: (i) Advance Payment: Ten (10) percent of the Contract Price shall be paid within thirty (30) days of signing of the Contract against a simple receipt and a bank guarantee for the equivalent amount and in the form provided in the bidding documents or another form acceptable to the Purchaser. (ii) On Delivery: Eighty (80) percent of the Contract Price shall be paid on receipt of the Goods and upon submission of the documents specified in GCC Clause 13. (iii) On Acceptance: The remaining ten (10) percent of the Contract Price shall be paid to the Supplier within thirty (30) days after the date of the acceptance certificate for the respective delivery issued by the Purchaser. |
| GCC 16.5 | The payment-delay period after which the Purchaser shall pay interest to the supplier shall be 45 days. The interest rate that shall be applied is 6 % |
| GCC 18.1 | A Performance Security “shall” be required 10 percentage of the Contract Price. |
| GCC 18.3 | If required, the Performance Security shall be in the form of : “a Demand Guarantee” If required, the Performance security shall be denominated in “The currencies of payment of the Contract, in accordance with their portions of the Contract Price” |
| GCC 18.4 | Discharge of the Performance Security shall take place: The percentage of the Performance security will be reduces up to 3% after all goods delivering, installation and training. |
| GCC 23.2 | The packing, marking and documentation within and outside the packages shall be: Information according to GCC 1.1 (i); GCC 1.1 (j) and GCC (o) |
| GCC 24.1 | The insurance coverage shall be as specified in the Incoterms. The Insurance shall be in an amount equal to 110% of the CIP price of Goods from “warehouse” to “warehouse” on “All Risks” basis, including War Risk and Strikes. |
| GCC 25.1 | “The Supplier is required under the Contract to transport the Goods to a specified place of final destination within the Purchaser’s country, defined as the Project Site, transport to such place of destination in the Purchaser’s country, including insurance and storage, as shall be specified in the Contract, shall be arranged by the Supplier, and related costs shall be included in the Contract Price” |
| GCC 25.2 | Incidental services to be provided are: Installation of all equipment. Within 25 days after delivery to the final destination of goods. Training, within 15 days after installation. No. of Persons and period of time - Lot 1 - 13 persons (doctors-5; X-ray technicians-7; nurse-1) to be trained for, during 5 days in “Vanadzor MC” CJSC. - Lot 2 - 21 persons (doctors-5; laboratory technicians-16) to be trained for Lot 2, during 5 days in “Vanadzor MC” CJSC. - Lot 3 - 28 persons (doctors-12; anesthesiology technicians-16) to be trained for Lot 3during 5 days in “Vanadzor MC” CJSC. |
| GCC 26.1 | The inspections and tests shall be: For Lots 1; 2 and 3 - The Supplier should conduct the following tests and verification supplied equipment: - Test ensuring that there are no physical damages to supplied, goods before delivery to the final destination. The Manufacturer’s factory testing certificate for goods supplied will be provided with delivery documents, according to SCC (in case of supply based on CIP terms). B. The following tests and inspections shall be carried out by the Supplier (or its designated representative) to the Final destination and conducted in the presence of the Purchaser, Beneficiary or their designated representative. - Verify the completeness of consumables and accessories and other spare parts (if applicable) of the equipments supplied. - Testing the operation of main function of the equipments supplied, after installation. C. The Supplier shall provide the Purchaser with a report of the results of above mentioned tests and with a report of completion of training. The names of the trained personal shall be detailed in the report. After that an appropriate delivery-acceptance act shall be issued and signed by the Purchaser, Beneficiary and Supplier. For Lot- 4 A. The Supplier should conduct the following tests and verification supplied equipment: - Test ensuring that there are no physical damages to supplied, goods (Medical Furniture), before delivery to the final destination. The Manufacturer’s factory testing certificate for goods supplied will be provided with delivery documents, according to SCC (in case of supply based on CIP terms). B. The following tests and inspections shall be carried out by the Supplier (or its designated representative) to the Final destination and conducted in the presence of the Purchaser, Beneficiary or their designated representative. - Verify the completeness of accessories and other spare parts (if applicable) of the Medical Furniture supplied. - Testing the operation of main function of the Medical Furniture supplied, after installation. C. The Supplier shall provide the Purchaser with a report of the results of above mentioned tests. After that an appropriate delivery-acceptance act shall be issued and signed by the Purchaser, Beneficiary and Supplier. Beneficiary is director or authorized representative from "Vanadzor MC" CJSC. |
| GCC 26.2 | The Inspections and tests shall be conducted at: “Vanadzor MC” CJSC Armenia, Lori Marz, Vanadzor Banaki St., 8a Build. |
| GCC 27.1 | The liquidated damage shall be: 0.5 % per week |
| GCC 27.1 | The maximum amount of liquidated damages shall be: 10 % of the Contract Price. |
| GCC 28.3 | The period of validity of the Warranty shall be: 2 (two) years For purposes of the Warranty, the place(s) of final destination(s) shall be: “Vanadzor MC” CJSC Armenia, Lori Marz, Vanadzor Banaki St., 8a Build.Sample provision. The warranty period shall be 2 (two) years from date of acceptance of the Goods, after installation. |
| GCC 28.5 | The period for repair or replacement shall be: - The quick respond period is 5 days, - For small and easy (small issues not impacting on the item functionality in general) defects repair or replacement period is 10 days, - For serious (issues impacting on the item functionality in general) defects repair or replacement period is 45 days (for Lots 1 and 3), - For serious (issues impacting on the item functionality in general) defects repair or replacement period is 30 days (for Lots 2 and 4), - For CT and X-Ray equipment (under Lot 1) is 3 months. |

Section X. Contract Forms

This Section contains forms which, once completed, will form part of the Contract. The forms for Performance Security and Advance Payment Security, when required, shall only be completed by the successful Bidder after contract award.

Table of Forms

Letter of Acceptance	220

1. Contract Agreement	221

2. Performance Security	223

3. Advance Payment Security	225

Letter of Acceptance

[letterhead paper of the Purchaser]

[date]

To:  [name and address of the Supplier]

Subject: Notification of Award Contract No.  . . . . . . . . . .

This is to notify you that your Bid dated . . . . [insert date] . . . .  for execution of the . . . . . . . . . .[insert name of the contract and identification number, as given in the SCC]. . . . . . . . . . for the Accepted Contract Amount of . . . . . . . . .[insert amount in numbers and words and name of currency], as corrected and modified in accordance with the Instructions to Bidders is hereby accepted by our Agency.

You are requested to furnish the Performance Security within 28 days in accordance with the Conditions of Contract, using for that purpose the of the Performance Security Form included in Section X, Contract Forms, of the Bidding Document.

Authorized Signature:

Name and Title of Signatory:

Name of Agency:

Attachment:  Contract Agreement

Contract Agreement

[The successful Bidder shall fill in this form in accordance with the instructions indicated]

THIS  AGREEMENT made

the [ insert:  number ] day of  [ insert:  month ], [ insert:  year ].

BETWEEN

(1)	[ insert complete name of Purchaser ], a [ insert description of type of legal entity, for example, an agency of the Ministry of .... of the Government of { insert name of  Country of Purchaser }, or corporation incorporated under the laws of { insert name of  Country of Purchaser } ] and having its principal place of business at [ insert address of Purchaser ] (hereinafter called “the Purchaser”), of the one part, and

(2)	[ insert name of Supplier], a corporation incorporated under the laws of [ insert:  country of Supplier] and having its principal place of business at [ insert:  address of Supplier ] (hereinafter called “the Supplier”), of the other part :

WHEREAS the Purchaser invited bids for certain Goods and ancillary services, viz., [insert brief description of Goods and Services] and has accepted a Bid by the Supplier for the supply of those Goods and Services

The Purchaser and the Supplier agree as follows:

1.	In this Agreement words and expressions shall have the same meanings as are respectively assigned to them in the Contract documents referred to.

2.	The following documents shall be deemed to form and be read and construed as part of this Agreement.  This Agreement shall prevail over all other contract documents.

- the Letter of Acceptance

- the Letter of Bid

- the Addenda Nos._____ (if any)

- Special Conditions of Contract

- General Conditions of Contract

- the Specification (including Schedule of Requirements and Technical Specifications)

- the completed Schedules (including Price Schedules)

- any other document listed in GCC as forming part of the Contract

3.	In consideration of the payments to be made by the Purchaser to the Supplier as specified in this Agreement,  the Supplier hereby covenants with the Purchaser to provide the Goods and Services and to remedy defects therein in conformity in all respects with the provisions of the Contract.

4.	The Purchaser hereby covenants to pay the Supplier in consideration of the provision of the Goods and Services and the remedying of defects therein, the Contract Price or such other sum as may become payable under the provisions of the Contract at the times and in the manner prescribed by the Contract.

IN WITNESS whereof the parties hereto have caused this Agreement to be executed in accordance with the laws of [insert the name of the Contract governing law country] on the day, month and year indicated above.

For and on behalf of the Purchaser

Signed:	[insert signature]

in the capacity of [ insert  title or other appropriate designation ]

in the presence of [insert identification of official witness]

For and on behalf of the Supplier

Signed:	[insert signature of authorized representative(s) of the Supplier]

in the capacity of [ insert  title or other appropriate designation ]

in the presence of [ insert identification of official witness]

Performance Security

Option 1: (Bank Guarantee)

[The bank, as requested by the successful Bidder, shall fill in this form in accordance with the instructions indicated]

[Guarantor letterhead or SWIFT identifier code]

Beneficiary:	[insert name and Address of Purchaser ]

Date:	_ [Insert date of issue]

PERFORMANCE GUARANTEE No.:	[Insert guarantee reference number]

Guarantor:  [Insert name and address of place of issue, unless indicated in the letterhead]

We have been informed that _ [insert name of Supplier, which in the case of a joint venture shall be the name of the joint venture] (hereinafter called "the Applicant") has entered into Contract No. [insert reference number of the contract] dated [insert date]with the Beneficiary, for the supply of _ [insert name of contract and brief description of Goods and related Services](hereinafter called "the Contract").

Furthermore, we understand that, according to the conditions of the Contract, a performance guarantee is required.

At the request of the Applicant, we as Guarantor, hereby irrevocably undertake to pay the Beneficiary any sum or sums not exceeding in total an amount of[insert amount in figures]
()[insert amount in words],1 such sum being payable in the types and proportions of currencies in which the Contract Price is payable, upon receipt by us of the Beneficiary’s complying demand supported by the Beneficiary’s statement, whether in the demand itself or in a separate signed document accompanying or identifying the demand, stating that the Applicant is in breach of its obligation(s) under the Contract, without the Beneficiary needing to prove or to show grounds for your demand or the sum specified therein.

This guarantee shall expire, no later than the …. Day of ……, 2… 2, and any demand for payment under it must be received by us at this office indicated above on or before that date.

This guarantee is subject to the Uniform Rules for Demand Guarantees (URDG) 2010 Revision, ICC Publication No. 758, except that the supporting statement under Article 15(a) is hereby excluded.

_____________________ 
[signature(s)]

Note:  All italicized text (including footnotes) is for use in preparing this form and shall be deleted from the final product.

Advance Payment Security

[Guarantor letterhead or SWIFT identifier code]

Beneficiary:[Insert name and Address of Purchaser]

Date:	[Insert date of issue]

ADVANCE PAYMENT GUARANTEE No.:	[Insert guarantee reference number]

Guarantor:  [Insert name and address of place of issue, unless indicated in the letterhead]

We have been informed that [insert name of Supplier, which in the case of a joint venture shall be the name of the joint venture] (hereinafter called “the Applicant”) has entered into Contract No. [insert reference number of the contract] dated [insert date] with the Beneficiary, for the execution of [insert name of contract and brief description of Goods and related Services] (hereinafter called "the Contract").

Furthermore, we understand that, according to the conditions of the Contract, an advance payment in the sum [insert amount in figures] () [insert amount in words] is to be made against an advance payment guarantee.

At the request of the Applicant, we as Guarantor, hereby irrevocably undertake to pay the Beneficiary any sum or sums not exceeding in total an amount of [insert amount in figures] 
() [insert amount in words]1 upon receipt by us of the Beneficiary’s complying demand supported by the Beneficiary’s statement, whether in the demand itself or in a separate signed document accompanying or identifying the demand, stating either that the Applicant:

    - has used the advance payment for purposes other than toward delivery of Goods; or

    - has failed to repay the advance payment in accordance with the Contract conditions, specifying the amount which the Applicant has failed to repay.

A demand under this guarantee may be presented as from the presentation to the Guarantor of a certificate from the Beneficiary’s bank stating that the advance payment referred to above has been credited to the Applicant on its account number [insert number] at  [insert name and address of Applicant’s bank].

This guarantee shall expire, at the latest, upon our receipt of a copy of the certificate indicating that goods have been delivered and  the services have been performed accordingly, or on the [insert day] day of [insert month], [insert year], whichever is earlier. Consequently, any demand for payment under this guarantee must be received by us at this office on or before that date.

This guarantee is subject to the Uniform Rules for Demand Guarantees (URDG) 2010 Revision, ICC Publication No.758,except that the supporting statement under Article 15(a) is hereby excluded.

.

____________________ 
[signature(s)]

Note:  All italicized text (including footnotes) is for use in preparing this form and shall be deleted from the final product.

Invitation for Bids

# REPUBLIC OF ARMENIA

# DISEASE PREVENTION AND CONTROL PROJECT

Credit No.: 5222

Contract Title: Medical, diagnostic equipment, Medical supplies and furniture for “Vanadzor MC” CJSC, comprising of 4 lots

Reference No.: CR4/ICB/B-G/013-16

1.	The Republic of Armenia has received financing from the World Bank toward the cost of the Disease prevention and control project, and intends to apply part of the proceeds toward payments under the contract for Medical, diagnostic equipment, Medical supplies and furniture for “Vanadzor MC” CJSC, comprising of 4 lots.  “Bidders may bid for one or several contracts, as further defined in the bidding document.  Bidders wishing to offer discounts in case they are awarded more than one contract will be allowed to do so, provided those discounts are included in the Letter of Bid”.

2. 	The Health project implementation unit, state agency now invites sealed bids from eligible bidders for Lot 1- Supply and Installation of Diagnostic and Imaging equipment for "Vanadzor MC” CJSC #CR4/ICB/B-G/013-16/1; Lot 2- Supply and Installation of Medical Laboratory equipment for "Vanadzor MC” CJSC #CR4/ICB/B-G/013-16/2; Lot 3 - Supply and Installation of Surgery and Hospital equipment for "Vanadzor MC” CJSC #CR4/ICB/B-G/013-16/3; Lot 4 -Supply and Installation of Medical Furniture for "Vanadzor MC” CJSC #CR4/ICB/B-G/013-16/4.

3. 	Bidding will be conducted through the International Competitive Bidding procedures as specified in the World Bank’s Guidelines: Procurement of Goods, Works and Non-Consulting Services under IBRD Loans and IDA Credits &Grants by World Bank Borrowers, January 2011, revised July 2014 (“Procurement Guidelines”), and is open to all eligible bidders as defined in the Procurement Guidelines. In addition, please refer to paragraphs 1.6 and 1.7 setting forth the World Bank’s policy on conflict of interest.

4. 	Interested eligible bidders may obtain further information from Health project implementation unit, state agency; e-mail procurement@healthpiu.am; info@healthpiu.am. The bidding documents will be available on the www.procurement.am (http://gnumner.am/am/category/129/1.html) website free of charge. However the HPIU SA will not be responsible for incorrect or incomplete downloading documents by the potential bidders, or missed amendments to those (if any).”

5. 	A complete set of bidding documents in English language may be purchased by interested eligible bidders upon the submission of a written application to the address below and upon payment of a nonrefundable fee of AMD 30,000 or equal in USD, in case of request for delivery the postal fee should be also reimbursed by the receiver. The method of payment will be bank transfer, details are below.

Central Treasury of Ministry of Finance, RA- 900000902123 for AMD transfers.

Central Treasury of Ministry of Finance, RA- 900000902156 for USD transfers.

6. 	Bids must be delivered to the address below on or before 01 March, 2017 10:00, local time. Electronic bidding will not be permitted. Late bids will be rejected. Bids will be publicly opened in the presence of the bidders, designated representatives and anyone who choose to attend at the address below on 10:00, local time on 01 March, 2017.

7. Bids must be accompanied by a “Bid Security” of:

Lot 1: USD 26,500

Lot 2: USD 1,600

Lot 3: USD 26,500

Lot 4: USD 6,600

8.	The address referred to above is:

Health project implementation unit, state agency

Attn: Mr. Nelson Zuloyan, director

Mr. Hovakim Podosyan, procurement specialist

Komitas 49/4 Avenue, 5th floor

Yerevan 0051, Republic of Armenia

Tel: +374 10 297535; +374 10 297537

Web site: www.healthpiu.am

Bidder to use as appropriate

The amount of the Bond shall be denominated in the currency of the Purchaser’s country or the equivalent amount in a freely convertible currency.

In this context, any action to influence the procurement process or contract execution for undue advantage is improper.

For the purpose of this sub-paragraph, “another party” refers to a public official acting in relation to the procurement process or contract execution. In this context, “public official” includes World Bank staff and employees of other organizations taking or reviewing procurement decisions.

For the purpose of this sub-paragraph, “party” refers to a public official; the terms “benefit” and “obligation” relate to the procurement process or contract execution; and the “act or omission” is intended to influence the procurement process or contract execution.

For the purpose of this sub-paragraph, “parties” refers to participants in the procurement process (including public officials) attempting either themselves, or through another person or entity not participating in the procurement or selection process, to simulate competition or to establish bid prices at artificial, non-competitive levels, or are privy to each other’s bid prices or other conditions.

For the purpose of this sub-paragraph, “party” refers to a participant in the procurement process or contract execution.

A firm or individual may be declared ineligible to be awarded a Bank financed contract upon: (i) completion of the Bank’s sanctions proceedings as per its sanctions procedures, including, inter alia, cross-debarment as agreed with other International Financial Institutions, including Multilateral Development Banks, and through the application the World Bank Group corporate administrative procurement sanctions procedures for fraud and corruption; and (ii) as a result of temporary suspension or early temporary suspension in connection with an ongoing sanctions proceeding. See footnote 14 and paragraph 8 of Appendix 1 of these Guidelines.

A nominated sub-contractor, consultant, manufacturer or supplier, or service provider (different names are used depending on the particular bidding document) is one which has either been: (i) included by the bidder in its pre-qualification application or bid because it brings specific and critical experience and know-how that allow the bidder to meet the qualification requirements for the particular bid; or (ii) appointed by the Borrower.

In this context, any action to influence the procurement process or contract execution for undue advantage is improper.

For the purpose of this sub-paragraph, “another party” refers to a public official acting in relation to the procurement process or contract execution. In this context, “public official” includes World Bank staff and employees of other organizations taking or reviewing procurement decisions.

For the purpose of this sub-paragraph, “party” refers to a public official; the terms “benefit” and “obligation” relate to the procurement process or contract execution; and the “act or omission” is intended to influence the procurement process or contract execution.

For the purpose of this sub-paragraph, “parties” refers to participants in the procurement process (including public officials) attempting either themselves, or through another person or entity not participating in the procurement or selection process, to simulate competition or to establish bid prices at artificial, non-competitive levels, or are privy to each other’s bid prices or other conditions.

For the purpose of this sub-paragraph, “party” refers to a participant in the procurement process or contract execution.

A firm or individual may be declared ineligible to be awarded a Bank financed contract upon: (i) completion of the Bank’s sanctions proceedings as per its sanctions procedures, including, inter alia, cross-debarment as agreed with other International Financial Institutions, including Multilateral Development Banks, and through the application the World Bank Group corporate administrative procurement sanctions procedures for fraud and corruption; and (ii) as a result of temporary suspension or early temporary suspension in connection with an ongoing sanctions proceeding. See footnote 14 and paragraph 8 of Appendix 1 of these Guidelines.

A nominated sub-contractor, consultant, manufacturer or supplier, or service provider (different names are used depending on the particular bidding document) is one which has either been: (i) included by the bidder in its pre-qualification application or bid because it brings specific and critical experience and know-how that allow the bidder to meet the qualification requirements for the particular bid; or (ii) appointed by the Borrower.

1	 The Guarantor shall insert an amount representing the percentage of the Accepted Contract Amount specified in the Letter of Acceptance, and denominated either in the currency(ies) of the Contract or a freely convertible currency acceptable to the Beneficiary.

2	Insert the date twenty-eight days after the expected completion dateas described in GC Clause 18.4. The Purchaser should note that in the event of an extension of this date for completion of the Contract, thePurchaser would need to request an extension of this guarantee from the Guarantor.  Such request must be in writing and must be made prior to the expiration date established in the guarantee. In preparing this guarantee, the Purchaser might consider adding the following text to the form, at the end of the penultimate paragraph:  “The Guarantor agrees to a one-time extension of this guarantee for a period not to exceed [six months][one year], in response to the Beneficiary’s written request for such extension, such request to be presented to the Guarantor before the expiry of the guarantee.”

1	The Guarantor shall insert an amount representing the amount of the advance payment and denominated either in the currency(ies) of the advance payment as specified in the Contract, or in a freely convertible currency acceptable to the Purchaser.